## OD600 reference point
| LUDOX-HS40 | H2O
Replicate 1 | 0.047 | 0.031 | Enter Abs600 absorbance measurements into blue cells
Replicate 2 | 0.042 | 0.03 | Gold cells are calculated
Replicate 3 | 0.041 | 0.029
Replicate 4 | 0.041 | 0.029
Arith. Mean | 0.04275 | 0.02975
Corrected Abs600 | 0.013 |  | Corrected value is particle-only contribution
Reference OD600 | 0.0425 |  | Reference value is for 100uL of LUDOX-HS40 in a well of a standard 96-well flat-bottom plate
OD600/Abs600 | 3.2692 |  | Corrected value = scaling factor * measured value

## Fluorescein standard curve
uM Fluorescein | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625 | 0.78125 | 0.39062 | 0.19531 | 0.097656 | 0.048828 | 0
Replicate 1 | 142265 | 91852 | 54301 | 29880 | 14142 | 9317 | 4899 | 3020 | 1221 | 728 | 256 | 22 | Enter fluorescence measurements into blue cells
Replicate 2 | 138310 | 91952 | 54105 | 29864 | 15335 | 7926 | 4157 | 1787 | 775 | 303 | 116 | 25 | Gold cells are calculated
Replicate 3 | 139051 | 91249 | 43601 | 47220 | 2456 | 5793 | 3719 | 1925 | 904 | 588 | 389 | 23
Replicate 4 | 139746 | 92102 | 54862 | 30209 | 14718 | 7984 | 3536 | 1490 | 603 | 318 | 115 | 26 | Values measured are fluorescence from 100uL of X uM fluorescein solution
Arith. Mean | 139843 | 91788.75 | 51717.25 | 34293.25 | 11662.75 | 7755 | 4077.75 | 2055.5 | 875.75 | 484.25 | 219 | 24 | Gain is 748 instead of 1000
Arith. Std.Dev. | 1717.8325 | 374.2133 | 5420.3346 | 8619.2999 | 6157.134 | 1457.2771 | 606.3351 | 668.125 | 261.1122 | 208.7013 | 131.2682 | 1.8257
Values should form a straight line on both linear and log scale
Slope should be 1:1
Common problems:
* Consistent pipetting error --> log graph is a straight line but not 1:1 slope
* Oversaturated detector --> low concentrations linear, but high concentrations saturate or fall
uM Fluorescein/a.u. | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625 | 0.78125 | 0.39062 | 0.19531 | 0.097656 | 0.048828
Mean um Fluorescein/a.u. | 0.00035754 | 0.00027236 | 0.0002417 | 0.00018225 | 0.00026795 | 0.00020148 | 0.00019159 | 0.00019004 | 0.00022302 | 0.00020166 | 0.00022296
Mean of med-high levels: |  | 0.00023315
Final scaling level determined from medium-high points likely to be less impacted by saturation or pipetting error
If needed, you can shift which points are used, but it is likely better to correct instrument settings and protocol.

## Raw Plate Reader Measurements
Raw Plate Readings |  | If you followed the recommended plate layout:
Copy fluorescence and Abs600 measurements from your plate reader into blue cells
They will automatically propagate into the correct locations in the Fluorescence Measurement Sheet
Fluorescence Raw Readings: | [9 empty] | Abs600 Raw Readings:
Hour 0: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 0: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 533 | 695 | 1726 | 988 | 490 | 719 | 523 | 492 | 435 | Colony 1, Replicate 1 | 0.521 | 0.478 | 0.486 | 0.593 | 0.533 | 0.578 | 0.557 | 0.573 | 0.615
Colony 1, Replicate 2 | 521 | 683 | 1681 | 992 | 485 | 745 | 533 | 506 | 306 | Colony 1, Replicate 2 | 0.552 | 0.579 | 0.496 | 0.588 | 0.595 | 0.553 | 0.477 | 0.534 | 0.675
Colony 1, Replicate 3 | 510 | 746 | 1674 | 1015 | 471 | 703 | 526 | 478 | 315 | Colony 1, Replicate 3 | 0.439 | 0.457 | 0.528 | 0.567 | 0.596 | 0.596 | 0.525 | 0.487 | 0.666
Colony 1, Replicate 4 | 523 | 706 | 1650 | 976 | 437 | 724 | 486 | 510 | 299 | Colony 1, Replicate 4 | 0.458 | 0.514 | 0.502 | 0.561 | 0.681 | 0.664 | 0.63 | 0.42 | 0.951
Colony 2, Replicate 1 | 769 | 600 | 1307 | 939 | 494 | 809 | 544 | 346 | 499 | Colony 2, Replicate 1 | 0.613 | 0.854 | 0.822 | 0.589 | 0.608 | 0.532 | 0.465 | 0.649 | 0.419
Colony 2, Replicate 2 | 685 | 705 | 1261 | 815 | 518 | 798 | 590 | 448 | 486 | Colony 2, Replicate 2 | 0.716 | 0.7 | 0.754 | 0.682 | 0.496 | 0.497 | 0.398 | 0.481 | 0.428
Colony 2, Replicate 3 | 528 | 687 | 462 | 488 | 518 | 799 | 480 | 321 | 368 | Colony 2, Replicate 3 | 0.895 | 0.727 | 0.972 | 0.947 | 0.539 | 0.492 | 0.5 | 0.89 | 0.59
Colony 2, Replicate 4 | 622 | 902 | 992 | 368 | 537 | 697 | 560 | 331 | 458 | Colony 2, Replicate 4 | 1.149 | 1.088 | 0.77 | 0.639 | 0.457 | 0.673 | 0.448 | 0.881 | 0.504
Hour 2: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 2: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 232 | 984 | 796 | 1220 | 253 | 941 | 481 | 251 | 239 | Colony 1, Replicate 1 | 0.567 | 0.519 | 0.385 | 0.518 | 0.515 | 0.397 | 0.508 | 0.555 | 0.287
Colony 1, Replicate 2 | 229 | 957 | 799 | 937 | 244 | 924 | 459 | 223 | 243 | Colony 1, Replicate 2 | 0.563 | 0.58 | 0.401 | 0.692 | 0.549 | 0.449 | 0.519 | 0.611 | 0.309
Colony 1, Replicate 3 | 232 | 953 | 753 | 1194 | 246 | 898 | 462 | 232 | 235 | Colony 1, Replicate 3 | 0.547 | 0.542 | 0.526 | 0.523 | 0.548 | 0.418 | 0.509 | 0.606 | 0.314
Colony 1, Replicate 4 | 235 | 998 | 743 | 1182 | 247 | 891 | 470 | 240 | 235 | Colony 1, Replicate 4 | 0.579 | 0.51 | 0.486 | 0.637 | 0.49 | 0.623 | 0.472 | 0.494 | 0.33
Colony 2, Replicate 1 | 1195 | 1173 | 1045 | 1216 | 237 | 959 | 532 | 247 | 239 | Colony 2, Replicate 1 | 0.489 | 0.577 | 0.545 | 0.494 | 0.638 | 0.575 | 0.639 | 0.551 | 0.314
Colony 2, Replicate 2 | 1158 | 1103 | 1025 | 1202 | 242 | 923 | 533 | 236 | 237 | Colony 2, Replicate 2 | 0.56 | 0.697 | 0.551 | 0.487 | 0.587 | 0.646 | 0.635 | 0.635 | 0.366
Colony 2, Replicate 3 | 1153 | 1120 | 1026 | 1166 | 228 | 896 | 511 | 235 | 233 | Colony 2, Replicate 3 | 0.472 | 0.564 | 0.527 | 0.612 | 0.484 | 0.665 | 0.489 | 0.503 | 0.366
Colony 2, Replicate 4 | 1146 | 1083 | 959 | 1158 | 205 | 940 | 521 | 227 | 242 | Colony 2, Replicate 4 | 0.516 | 0.65 | 0.408 | 0.633 | 0.571 | 0.416 | 0.668 | 0.678 | 0.361
0.33087
Hour 4: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 4: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 243 | 1533 | 875 | 2083 | 245 | 1717 | 531 | 250 | 228 | Colony 1, Replicate 1 | 0.704 | 0.625 | 0.329 | 0.651 | 0.609 | 0.468 | 0.633 | 0.669 | 0.257
Colony 1, Replicate 2 | 234 | 1455 | 878 | 2063 | 264 | 1623 | 543 | 230 | 234 | Colony 1, Replicate 2 | 0.74 | 0.596 | 0.323 | 0.641 | 0.625 | 0.445 | 0.671 | 0.668 | 0.329
Colony 1, Replicate 3 | 246 | 1500 | 883 | 2081 | 260 | 1643 | 533 | 233 | 235 | Colony 1, Replicate 3 | 0.706 | 0.603 | 0.328 | 0.615 | 0.641 | 0.469 | 0.637 | 0.637 | 0.279
Colony 1, Replicate 4 | 240 | 1474 | 884 | 1999 | 253 | 1650 | 529 | 242 | 249 | Colony 1, Replicate 4 | 0.68 | 0.598 | 0.336 | 0.605 | 0.606 | 0.499 | 0.604 | 0.623 | 0.3
Colony 2, Replicate 1 | 1845 | 1701 | 1294 | 2019 | 251 | 1756 | 611 | 243 | 241 | Colony 2, Replicate 1 | 0.615 | 0.612 | 0.353 | 0.617 | 0.623 | 0.509 | 0.612 | 0.623 | 0.279
Colony 2, Replicate 2 | 1821 | 1680 | 1316 | 2015 | 254 | 1548 | 626 | 236 | 253 | Colony 2, Replicate 2 | 0.657 | 0.596 | 0.385 | 0.6 | 0.611 | 0.716 | 0.622 | 0.632 | 0.29
Colony 2, Replicate 3 | 1797 | 1747 | 1223 | 1914 | 252 | 1724 | 604 | 232 | 242 | Colony 2, Replicate 3 | 0.596 | 0.638 | 0.332 | 0.583 | 0.672 | 0.536 | 0.616 | 0.753 | 0.483
Colony 2, Replicate 4 | 1759 | 1667 | 1223 | 1968 | 249 | 1662 | 610 | 227 | 249 | Colony 2, Replicate 4 | 0.629 | 0.592 | 0.555 | 0.704 | 0.644 | 0.466 | 0.617 | 0.767 | 0.233
Hour 6: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 6: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 238 | 1940 | 1101 | 3279 | 259 | 2687 | 557 | 249 | 232 | Colony 1, Replicate 1 | 0.827 | 0.815 | 0.487 | 0.816 | 0.821 | 0.72 | 0.852 | 0.89 | 0.461
Colony 1, Replicate 2 | 237 | 1933 | 1098 | 3258 | 252 | 2715 | 561 | 235 | 236 | Colony 1, Replicate 2 | 0.84 | 0.81 | 0.486 | 0.857 | 0.836 | 0.744 | 0.858 | 0.87 | 0.394
Colony 1, Replicate 3 | 228 | 1955 | 1102 | 3375 | 261 | 2723 | 563 | 235 | 235 | Colony 1, Replicate 3 | 0.847 | 0.816 | 0.462 | 0.841 | 0.841 | 0.753 | 0.823 | 0.878 | 0.357
Colony 1, Replicate 4 | 232 | 1952 | 1085 | 3331 | 252 | 2720 | 592 | 240 | 219 | Colony 1, Replicate 4 | 0.864 | 0.816 | 0.467 | 0.851 | 0.813 | 0.708 | 0.829 | 0.852 | 0.522
Colony 2, Replicate 1 | 2311 | 2155 | 1823 | 3224 | 254 | 2490 | 641 | 229 | 245 | Colony 2, Replicate 1 | 0.794 | 0.799 | 0.522 | 0.808 | 0.837 | 0.713 | 0.841 | 0.853 | 0.33
Colony 2, Replicate 2 | 2315 | 2198 | 1896 | 3314 | 258 | 2549 | 658 | 237 | 242 | Colony 2, Replicate 2 | 0.772 | 0.814 | 0.528 | 0.827 | 0.844 | 0.688 | 0.824 | 0.813 | 0.163
Colony 2, Replicate 3 | 2334 | 2148 | 1818 | 3325 | 257 | 2533 | 650 | 239 | 236 | Colony 2, Replicate 3 | 0.789 | 0.792 | 0.619 | 0.827 | 0.846 | 0.734 | 0.807 | 0.876 | 0.31
Colony 2, Replicate 4 | 2362 | 2059 | 1899 | 3297 | 258 | 2524 | 647 | 234 | 237 | Colony 2, Replicate 4 | 0.793 | 0.784 | 0.56 | 0.8 | 0.854 | 0.749 | 0.862 | 0.843 | 0.358
Plate pattern:
| A1 | A2 | A3 | A4 | A5 | A6 | A7 | A8 | A9
| B1 | B2 | B3 | B4 | B5 | B6 | B7 | B8 | B9
| C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9
| D1 | D2 | D3 | D4 | D5 | D6 | D7 | D8 | D9
| E1 | E2 | E3 | E4 | E5 | E6 | E7 | E8 | E9
| F1 | F2 | F3 | F4 | F5 | F6 | F7 | F8 | F9
| G1 | G2 | G3 | G4 | G5 | G6 | G7 | G8 | G9
| H1 | H2 | H3 | H4 | H5 | H6 | H7 | H8 | H9

## Fluorescence Measurement
Unit Scaling Factors: | These are imported from the prior two sheets | [5 empty] | Enter fluorescence and Abs600 measurements into blue cells on "Raw Plate Reader Measurements"
OD600/Abs600 | 3.2692 | [5 empty] | They will be copied into the green cells on this sheet.
uM Fluorescein/a.u. | 0.00023315 | [5 empty] | Gold cells are calculated
If you have more replicates, unhide the extra columns
Experimental Values: | Raw Abs600 | [5 empty] | Raw Fluorescence
Sample set: | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6
Blank media | 0.287 | 0.309 | 0.314 | 0.33 |  |  | 435 | 306 | 315 | 299
Blank mean: | 0.31 | [5 empty] | 338.75 | [5 empty] | OD - Background | [5 empty] | Fluorescence - Background | [5 empty] | uM Fluorescein / OD600 | [5 empty] | Summary Statistics |  |  |  | Ln uM FITC / OD600
Hour 0: | [12 empty] | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Arith. Mean | Arith. Std.Dev. | Geo. Mean | Geo. Std. Dev. | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6
Negative Control (Colony 1) | 0.521 | 0.552 | 0.439 | 0.458 |  |  | 533 | 521 | 510 | 523 |  |  | 0.211 | 0.242 | 0.129 | 0.148 | --- | --- | 194.25 | 182.25 | 171.25 | 184.25 | --- | --- | 0.065655 | 0.053708 | 0.094674 | 0.088784 | --- | --- | 0.075705 | 0.019285 | 0.073785 | 1.3036 | -2.7233 | -2.9242 | -2.3573 | -2.4216 | --- | ---
Negative Control (Colony 2) | 0.613 | 0.716 | 0.895 | 1.149 |  |  | 769 | 685 | 528 | 622 |  |  | 0.303 | 0.406 | 0.585 | 0.839 | --- | --- | 430.25 | 346.25 | 189.25 | 283.25 | --- | --- | 0.10127 | 0.060821 | 0.023071 | 0.024077 | --- | --- | 0.052309 | 0.037064 | 0.043008 | 2.0652 | -2.29 | -2.7998 | -3.7692 | -3.7265 | --- | ---
Positive Control (Colony 1) | 0.478 | 0.579 | 0.457 | 0.514 |  |  | 695 | 683 | 746 | 706 |  |  | 0.168 | 0.269 | 0.147 | 0.204 | --- | --- | 356.25 | 344.25 | 407.25 | 367.25 | --- | --- | 0.15123 | 0.091266 | 0.19757 | 0.12839 | --- | --- | 0.14211 | 0.044471 | 0.13679 | 1.3813 | -1.889 | -2.394 | -1.6216 | -2.0527 | --- | ---
Positive Control (Colony 2) | 0.854 | 0.7 | 0.727 | 1.088 |  |  | 600 | 705 | 687 | 902 |  |  | 0.544 | 0.39 | 0.417 | 0.778 | --- | --- | 261.25 | 366.25 | 348.25 | 563.25 | --- | --- | 0.034249 | 0.066973 | 0.059558 | 0.051631 | --- | --- | 0.053103 | 0.014044 | 0.051535 | 1.3397 | -3.3741 | -2.7035 | -2.8208 | -2.9636 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.486 | 0.496 | 0.528 | 0.502 |  |  | 1726 | 1681 | 1674 | 1650 |  |  | 0.176 | 0.186 | 0.218 | 0.192 | --- | --- | 1387.25 | 1342.25 | 1335.25 | 1311.25 | --- | --- | 0.56212 | 0.51465 | 0.43681 | 0.48705 | --- | --- | 0.50016 | 0.05239 | 0.49808 | 1.1113 | -0.57604 | -0.66428 | -0.82825 | -0.71939 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.822 | 0.754 | 0.972 | 0.77 |  |  | 1307 | 1261 | 462 | 992 |  |  | 0.512 | 0.444 | 0.662 | 0.46 | --- | --- | 968.25 | 922.25 | 123.25 | 653.25 | --- | --- | 0.13487 | 0.14813 | 0.013278 | 0.10128 | --- | --- | 0.099389 | 0.0607 | 0.071994 | 3.1223 | -2.0035 | -1.9096 | -4.3217 | -2.2899 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.593 | 0.588 | 0.567 | 0.561 |  |  | 988 | 992 | 1015 | 976 |  |  | 0.283 | 0.278 | 0.257 | 0.251 | --- | --- | 649.25 | 653.25 | 676.25 | 637.25 | --- | --- | 0.16361 | 0.16758 | 0.18766 | 0.18106 | --- | --- | 0.17498 | 0.011279 | 0.17471 | 1.0665 | -1.8103 | -1.7863 | -1.6731 | -1.7089 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.589 | 0.682 | 0.947 | 0.639 |  |  | 939 | 815 | 488 | 368 |  |  | 0.279 | 0.372 | 0.637 | 0.329 | --- | --- | 600.25 | 476.25 | 149.25 | 29.25 | --- | --- | 0.15343 | 0.091302 | 0.016709 | 0.0063404 | --- | --- | 0.066946 | 0.068968 | 0.034904 | 4.3918 | -1.8745 | -2.3936 | -4.0918 | -5.0608 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.533 | 0.595 | 0.596 | 0.681 |  |  | 490 | 485 | 471 | 437 |  |  | 0.223 | 0.285 | 0.286 | 0.371 | --- | --- | 151.25 | 146.25 | 132.25 | 98.25 | --- | --- | 0.04837 | 0.036596 | 0.032977 | 0.018886 | --- | --- | 0.034208 | 0.012146 | 0.032404 | 1.4837 | -3.0289 | -3.3078 | -3.4119 | -3.9693 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.608 | 0.496 | 0.539 | 0.457 |  |  | 494 | 518 | 518 | 537 |  |  | 0.298 | 0.186 | 0.229 | 0.147 | --- | --- | 155.25 | 179.25 | 179.25 | 198.25 | --- | --- | 0.037154 | 0.068728 | 0.055823 | 0.09618 | --- | --- | 0.064471 | 0.024796 | 0.06085 | 1.4889 | -3.2927 | -2.6776 | -2.8856 | -2.3415 | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.578 | 0.553 | 0.596 | 0.664 |  |  | 719 | 745 | 703 | 724 |  |  | 0.268 | 0.243 | 0.286 | 0.354 | --- | --- | 380.25 | 406.25 | 364.25 | 385.25 | --- | --- | 0.10119 | 0.11923 | 0.090828 | 0.077612 | --- | --- | 0.097213 | 0.017563 | 0.096031 | 1.1981 | -2.2908 | -2.1267 | -2.3988 | -2.556 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.532 | 0.497 | 0.492 | 0.673 |  |  | 809 | 798 | 799 | 697 |  |  | 0.222 | 0.187 | 0.182 | 0.363 | --- | --- | 470.25 | 459.25 | 460.25 | 358.25 | --- | --- | 0.15106 | 0.17514 | 0.18035 | 0.070383 | --- | --- | 0.14423 | 0.05086 | 0.13537 | 1.5572 | -1.89 | -1.7421 | -1.7129 | -2.6538 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.557 | 0.477 | 0.525 | 0.63 |  |  | 523 | 533 | 526 | 486 |  |  | 0.247 | 0.167 | 0.215 | 0.32 | --- | --- | 184.25 | 194.25 | 187.25 | 147.25 | --- | --- | 0.053198 | 0.082953 | 0.062111 | 0.032817 | --- | --- | 0.05777 | 0.020789 | 0.054764 | 1.4739 | -2.9337 | -2.4895 | -2.7788 | -3.4168 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.465 | 0.398 | 0.5 | 0.448 |  |  | 544 | 590 | 480 | 560 |  |  | 0.155 | 0.088 | 0.19 | 0.138 | --- | --- | 205.25 | 251.25 | 141.25 | 221.25 | --- | --- | 0.094436 | 0.20362 | 0.053018 | 0.11434 | --- | --- | 0.11635 | 0.063536 | 0.10391 | 1.7417 | -2.3598 | -1.5915 | -2.9371 | -2.1686 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.573 | 0.534 | 0.487 | 0.42 |  |  | 492 | 506 | 478 | 510 |  |  | 0.263 | 0.224 | 0.177 | 0.11 | --- | --- | 153.25 | 167.25 | 139.25 | 171.25 | --- | --- | 0.041556 | 0.053248 | 0.056106 | 0.11103 | --- | --- | 0.065484 | 0.031007 | 0.060932 | 1.5233 | -3.1807 | -2.9328 | -2.8805 | -2.198 | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.649 | 0.481 | 0.89 | 0.881 |  |  | 346 | 448 | 321 | 331 |  |  | 0.339 | 0.171 | 0.58 | 0.571 | --- | --- | 7.25 | 109.25 | -17.75 | -7.75 | --- | --- | 0.0015252 | 0.045563 | -0.0021825 | -0.00096795 | --- | --- | 0.010984 | 0.023104 | #NUM! | #NUM! | -6.4856 | -3.0887 | #NUM! | #NUM! | --- | ---
Hour 2:
Negative Control (Colony 1) | 0.567 | 0.563 | 0.547 | 0.579 |  |  | 232 | 229 | 232 | 235 |  |  | 0.257 | 0.253 | 0.237 | 0.269 | --- | --- | -106.75 | -109.75 | -106.75 | -103.75 | --- | --- | -0.029623 | -0.030937 | -0.032122 | -0.027506 | --- | --- | -0.030047 | 0.0019779 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Negative Control (Colony 2) | 0.489 | 0.56 | 0.472 | 0.516 |  |  | 1195 | 1158 | 1153 | 1146 |  |  | 0.179 | 0.25 | 0.162 | 0.206 | --- | --- | 856.25 | 819.25 | 814.25 | 807.25 | --- | --- | 0.34114 | 0.2337 | 0.35845 | 0.27947 | --- | --- | 0.30319 | 0.057404 | 0.29894 | 1.2168 | -1.0755 | -1.4537 | -1.026 | -1.2749 | --- | ---
Positive Control (Colony 1) | 0.519 | 0.58 | 0.542 | 0.51 |  |  | 984 | 957 | 953 | 998 |  |  | 0.209 | 0.27 | 0.232 | 0.2 | --- | --- | 645.25 | 618.25 | 614.25 | 659.25 | --- | --- | 0.22018 | 0.1633 | 0.18882 | 0.23508 | --- | --- | 0.20184 | 0.032123 | 0.19987 | 1.1771 | -1.5133 | -1.8122 | -1.667 | -1.4478 | --- | ---
Positive Control (Colony 2) | 0.577 | 0.697 | 0.564 | 0.65 |  |  | 1173 | 1103 | 1120 | 1083 |  |  | 0.267 | 0.387 | 0.254 | 0.34 | --- | --- | 834.25 | 764.25 | 781.25 | 744.25 | --- | --- | 0.22283 | 0.14084 | 0.21935 | 0.15611 | --- | --- | 0.18478 | 0.042412 | 0.18106 | 1.2643 | -1.5013 | -1.9602 | -1.5171 | -1.8572 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.385 | 0.401 | 0.526 | 0.486 |  |  | 796 | 799 | 753 | 743 |  |  | 0.075 | 0.091 | 0.216 | 0.176 | --- | --- | 457.25 | 460.25 | 414.25 | 404.25 | --- | --- | 0.43479 | 0.3607 | 0.13677 | 0.1638 | --- | --- | 0.27402 | 0.14645 | 0.24346 | 1.7712 | -0.83289 | -1.0197 | -1.9894 | -1.8091 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.545 | 0.551 | 0.527 | 0.408 |  |  | 1045 | 1025 | 1026 | 959 |  |  | 0.235 | 0.241 | 0.217 | 0.098 | --- | --- | 706.25 | 686.25 | 687.25 | 620.25 | --- | --- | 0.21433 | 0.20307 | 0.22586 | 0.45137 | --- | --- | 0.27366 | 0.11884 | 0.25809 | 1.4552 | -1.5402 | -1.5942 | -1.4878 | -0.79548 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.518 | 0.692 | 0.523 | 0.637 |  |  | 1220 | 937 | 1194 | 1182 |  |  | 0.208 | 0.382 | 0.213 | 0.327 | --- | --- | 881.25 | 598.25 | 855.25 | 843.25 | --- | --- | 0.30215 | 0.11169 | 0.28635 | 0.18391 | --- | --- | 0.22102 | 0.08978 | 0.20532 | 1.5886 | -1.1968 | -2.192 | -1.2505 | -1.6933 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.494 | 0.487 | 0.612 | 0.633 |  |  | 1216 | 1202 | 1166 | 1158 |  |  | 0.184 | 0.177 | 0.302 | 0.323 | --- | --- | 877.25 | 863.25 | 827.25 | 819.25 | --- | --- | 0.34001 | 0.34782 | 0.19535 | 0.18088 | --- | --- | 0.26602 | 0.090199 | 0.25425 | 1.4194 | -1.0788 | -1.0561 | -1.633 | -1.7099 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.515 | 0.549 | 0.548 | 0.49 |  |  | 253 | 244 | 246 | 247 |  |  | 0.205 | 0.239 | 0.238 | 0.18 | --- | --- | -85.75 | -94.75 | -92.75 | -91.75 | --- | --- | -0.029831 | -0.028273 | -0.027792 | -0.036351 | --- | --- | -0.030562 | 0.0039566 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.638 | 0.587 | 0.484 | 0.571 |  |  | 237 | 242 | 228 | 205 |  |  | 0.328 | 0.277 | 0.174 | 0.261 | --- | --- | -101.75 | -96.75 | -110.75 | -133.75 | --- | --- | -0.022123 | -0.024909 | -0.045392 | -0.036546 | --- | --- | -0.032243 | 0.010764 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.397 | 0.449 | 0.418 | 0.623 |  |  | 941 | 924 | 898 | 891 |  |  | 0.087 | 0.139 | 0.108 | 0.313 | --- | --- | 602.25 | 585.25 | 559.25 | 552.25 | --- | --- | 0.49368 | 0.30027 | 0.36929 | 0.12583 | --- | --- | 0.32227 | 0.15348 | 0.28809 | 1.8016 | -0.70587 | -1.2031 | -0.99617 | -2.0728 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.575 | 0.646 | 0.665 | 0.416 |  |  | 959 | 923 | 896 | 940 |  |  | 0.265 | 0.336 | 0.355 | 0.106 | --- | --- | 620.25 | 584.25 | 557.25 | 601.25 | --- | --- | 0.16692 | 0.12401 | 0.11195 | 0.40452 | --- | --- | 0.20185 | 0.13716 | 0.17497 | 1.7929 | -1.7902 | -2.0874 | -2.1897 | -0.90506 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.508 | 0.519 | 0.509 | 0.472 |  |  | 481 | 459 | 462 | 470 |  |  | 0.198 | 0.209 | 0.199 | 0.162 | --- | --- | 142.25 | 120.25 | 123.25 | 131.25 | --- | --- | 0.051236 | 0.041032 | 0.044169 | 0.057779 | --- | --- | 0.048554 | 0.0074855 | 0.048128 | 1.1652 | -2.9713 | -3.1934 | -3.1197 | -2.8511 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.639 | 0.635 | 0.489 | 0.668 |  |  | 532 | 533 | 511 | 521 |  |  | 0.329 | 0.325 | 0.179 | 0.358 | --- | --- | 193.25 | 194.25 | 172.25 | 182.25 | --- | --- | 0.04189 | 0.042625 | 0.068627 | 0.036306 | --- | --- | 0.047362 | 0.014455 | 0.045926 | 1.3195 | -3.1727 | -3.1553 | -2.6791 | -3.3158 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.555 | 0.611 | 0.606 | 0.494 |  |  | 251 | 223 | 232 | 240 |  |  | 0.245 | 0.301 | 0.296 | 0.184 | --- | --- | -87.75 | -115.75 | -106.75 | -98.75 | --- | --- | -0.025543 | -0.027425 | -0.02572 | -0.038274 | --- | --- | -0.02924 | 0.0060821 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.551 | 0.635 | 0.503 | 0.678 |  |  | 247 | 236 | 235 | 227 |  |  | 0.241 | 0.325 | 0.193 | 0.368 | --- | --- | -91.75 | -102.75 | -103.75 | -111.75 | --- | --- | -0.02715 | -0.022547 | -0.038337 | -0.021656 | --- | --- | -0.027423 | 0.0076642 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Hour 4:
Negative Control (Colony 1) | 0.704 | 0.74 | 0.706 | 0.68 |  |  | 243 | 234 | 246 | 240 |  |  | 0.394 | 0.43 | 0.396 | 0.37 | --- | --- | -95.75 | -104.75 | -92.75 | -98.75 | --- | --- | -0.017331 | -0.017373 | -0.016703 | -0.019034 | --- | --- | -0.01761 | 0.0009971 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Negative Control (Colony 2) | 0.615 | 0.657 | 0.596 | 0.629 |  |  | 1845 | 1821 | 1797 | 1759 |  |  | 0.305 | 0.347 | 0.286 | 0.319 | --- | --- | 1506.25 | 1482.25 | 1458.25 | 1420.25 | --- | --- | 0.3522 | 0.30464 | 0.36363 | 0.31751 | --- | --- | 0.33449 | 0.02794 | 0.33361 | 1.0874 | -1.0436 | -1.1886 | -1.0116 | -1.1472 | --- | ---
Positive Control (Colony 1) | 0.625 | 0.596 | 0.603 | 0.598 |  |  | 1533 | 1455 | 1500 | 1474 |  |  | 0.315 | 0.286 | 0.293 | 0.288 | --- | --- | 1194.25 | 1116.25 | 1161.25 | 1135.25 | --- | --- | 0.27038 | 0.27834 | 0.28265 | 0.28112 | --- | --- | 0.27812 | 0.0054608 | 0.27808 | 1.02 | -1.3079 | -1.2789 | -1.2636 | -1.269 | --- | ---
Positive Control (Colony 2) | 0.612 | 0.596 | 0.638 | 0.592 |  |  | 1701 | 1680 | 1747 | 1667 |  |  | 0.302 | 0.286 | 0.328 | 0.282 | --- | --- | 1362.25 | 1341.25 | 1408.25 | 1328.25 | --- | --- | 0.32169 | 0.33445 | 0.30619 | 0.33591 | --- | --- | 0.32456 | 0.01381 | 0.32434 | 1.044 | -1.1342 | -1.0953 | -1.1835 | -1.0909 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.329 | 0.323 | 0.328 | 0.336 |  |  | 875 | 878 | 883 | 884 |  |  | 0.019 | 0.013 | 0.018 | 0.026 | --- | --- | 536.25 | 539.25 | 544.25 | 545.25 | --- | --- | 2.0128 | 2.9582 | 2.1563 | 1.4956 | --- | --- | 2.1557 | 0.6056 | 2.0933 | 1.323 | 0.69953 | 1.0846 | 0.7684 | 0.40252 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.353 | 0.385 | 0.332 | 0.555 |  |  | 1294 | 1316 | 1223 | 1223 |  |  | 0.043 | 0.075 | 0.022 | 0.245 | --- | --- | 955.25 | 977.25 | 884.25 | 884.25 | --- | --- | 1.5843 | 0.92925 | 2.8664 | 0.25739 | --- | --- | 1.4093 | 1.1122 | 1.0209 | 2.7936 | 0.46014 | -0.073378 | 1.0531 | -1.3572 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.651 | 0.641 | 0.615 | 0.605 |  |  | 2083 | 2063 | 2081 | 1999 |  |  | 0.341 | 0.331 | 0.305 | 0.295 | --- | --- | 1744.25 | 1724.25 | 1742.25 | 1660.25 | --- | --- | 0.36479 | 0.3715 | 0.40738 | 0.40136 | --- | --- | 0.38626 | 0.021237 | 0.38582 | 1.0566 | -1.0084 | -0.9902 | -0.89801 | -0.91288 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.617 | 0.6 | 0.583 | 0.704 |  |  | 2019 | 2015 | 1914 | 1968 |  |  | 0.307 | 0.29 | 0.273 | 0.394 | --- | --- | 1680.25 | 1676.25 | 1575.25 | 1629.25 | --- | --- | 0.39032 | 0.41222 | 0.4115 | 0.2949 | --- | --- | 0.37724 | 0.055822 | 0.37381 | 1.1736 | -0.94078 | -0.8862 | -0.88793 | -1.2211 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.609 | 0.625 | 0.641 | 0.606 |  |  | 245 | 264 | 260 | 253 |  |  | 0.299 | 0.315 | 0.331 | 0.296 | --- | --- | -93.75 | -74.75 | -78.75 | -85.75 | --- | --- | -0.022361 | -0.016923 | -0.016967 | -0.02066 | --- | --- | -0.019228 | 0.0027256 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.623 | 0.611 | 0.672 | 0.644 |  |  | 251 | 254 | 252 | 249 |  |  | 0.313 | 0.301 | 0.362 | 0.334 | --- | --- | -87.75 | -84.75 | -86.75 | -89.75 | --- | --- | -0.019994 | -0.02008 | -0.01709 | -0.019164 | --- | --- | -0.019082 | 0.0013905 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.468 | 0.445 | 0.469 | 0.499 |  |  | 1717 | 1623 | 1643 | 1650 |  |  | 0.158 | 0.135 | 0.159 | 0.189 | --- | --- | 1378.25 | 1284.25 | 1304.25 | 1311.25 | --- | --- | 0.6221 | 0.67843 | 0.58499 | 0.49478 | --- | --- | 0.59508 | 0.077112 | 0.5912 | 1.1426 | -0.47466 | -0.38798 | -0.53615 | -0.70364 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.509 | 0.716 | 0.536 | 0.466 |  |  | 1756 | 1548 | 1724 | 1662 |  |  | 0.199 | 0.406 | 0.226 | 0.156 | --- | --- | 1417.25 | 1209.25 | 1385.25 | 1323.25 | --- | --- | 0.5079 | 0.21241 | 0.43713 | 0.60493 | --- | --- | 0.44059 | 0.16695 | 0.41098 | 1.5834 | -0.67746 | -1.5492 | -0.82753 | -0.50264 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.633 | 0.671 | 0.637 | 0.604 |  |  | 531 | 543 | 533 | 529 |  |  | 0.323 | 0.361 | 0.327 | 0.294 | --- | --- | 192.25 | 204.25 | 194.25 | 190.25 | --- | --- | 0.042447 | 0.04035 | 0.042364 | 0.046149 | --- | --- | 0.042828 | 0.0024174 | 0.042777 | 1.0573 | -3.1595 | -3.2102 | -3.1614 | -3.0759 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.612 | 0.622 | 0.616 | 0.617 |  |  | 611 | 626 | 604 | 610 |  |  | 0.302 | 0.312 | 0.306 | 0.307 | --- | --- | 272.25 | 287.25 | 265.25 | 271.25 | --- | --- | 0.064291 | 0.065659 | 0.061819 | 0.063011 | --- | --- | 0.063695 | 0.0016531 | 0.063679 | 1.0263 | -2.7443 | -2.7233 | -2.7835 | -2.7644 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.669 | 0.668 | 0.637 | 0.623 |  |  | 250 | 230 | 233 | 242 |  |  | 0.359 | 0.358 | 0.327 | 0.313 | --- | --- | -88.75 | -108.75 | -105.75 | -96.75 | --- | --- | -0.01763 | -0.021664 | -0.023063 | -0.022044 | --- | --- | -0.0211 | 0.0023876 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.623 | 0.632 | 0.753 | 0.767 |  |  | 243 | 236 | 232 | 227 |  |  | 0.313 | 0.322 | 0.443 | 0.457 | --- | --- | -95.75 | -102.75 | -106.75 | -111.75 | --- | --- | -0.021816 | -0.022757 | -0.017185 | -0.017439 | --- | --- | -0.019799 | 0.0028995 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Hour 6:
Negative Control (Colony 1) | 0.827 | 0.84 | 0.847 | 0.864 |  |  | 238 | 237 | 228 | 232 |  |  | 0.517 | 0.53 | 0.537 | 0.554 | --- | --- | -100.75 | -101.75 | -110.75 | -106.75 | --- | --- | -0.013898 | -0.013691 | -0.014708 | -0.013742 | --- | --- | -0.01401 | 0.00047379 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Negative Control (Colony 2) | 0.794 | 0.772 | 0.789 | 0.793 |  |  | 2311 | 2315 | 2334 | 2362 |  |  | 0.484 | 0.462 | 0.479 | 0.483 | --- | --- | 1972.25 | 1976.25 | 1995.25 | 2023.25 | --- | --- | 0.29061 | 0.30506 | 0.29706 | 0.29874 | --- | --- | 0.29787 | 0.0059411 | 0.29782 | 1.0202 | -1.2358 | -1.1872 | -1.2138 | -1.2082 | --- | ---
Positive Control (Colony 1) | 0.815 | 0.81 | 0.816 | 0.816 |  |  | 1940 | 1933 | 1955 | 1952 |  |  | 0.505 | 0.5 | 0.506 | 0.506 | --- | --- | 1601.25 | 1594.25 | 1616.25 | 1613.25 | --- | --- | 0.22613 | 0.22739 | 0.2278 | 0.22737 | --- | --- | 0.22717 | 0.00072259 | 0.22717 | 1.0032 | -1.4867 | -1.4811 | -1.4793 | -1.4812 | --- | ---
Positive Control (Colony 2) | 0.799 | 0.814 | 0.792 | 0.784 |  |  | 2155 | 2198 | 2148 | 2059 |  |  | 0.489 | 0.504 | 0.482 | 0.474 | --- | --- | 1816.25 | 1859.25 | 1809.25 | 1720.25 | --- | --- | 0.26488 | 0.26308 | 0.26769 | 0.25882 | --- | --- | 0.26362 | 0.0037196 | 0.2636 | 1.0142 | -1.3285 | -1.3353 | -1.3179 | -1.3516 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.487 | 0.486 | 0.462 | 0.467 |  |  | 1101 | 1098 | 1102 | 1085 |  |  | 0.177 | 0.176 | 0.152 | 0.157 | --- | --- | 762.25 | 759.25 | 763.25 | 746.25 | --- | --- | 0.30712 | 0.30765 | 0.35811 | 0.33898 | --- | --- | 0.32796 | 0.025012 | 0.32726 | 1.0787 | -1.1805 | -1.1788 | -1.0269 | -1.0818 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.522 | 0.528 | 0.619 | 0.56 |  |  | 1823 | 1896 | 1818 | 1899 |  |  | 0.212 | 0.218 | 0.309 | 0.25 | --- | --- | 1484.25 | 1557.25 | 1479.25 | 1560.25 | --- | --- | 0.4993 | 0.50944 | 0.34141 | 0.44508 | --- | --- | 0.44881 | 0.076972 | 0.44339 | 1.2022 | -0.69455 | -0.67445 | -1.0747 | -0.80949 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.816 | 0.857 | 0.841 | 0.851 |  |  | 3279 | 3258 | 3375 | 3331 |  |  | 0.506 | 0.547 | 0.531 | 0.541 | --- | --- | 2940.25 | 2919.25 | 3036.25 | 2992.25 | --- | --- | 0.4144 | 0.3806 | 0.40778 | 0.39445 | --- | --- | 0.39931 | 0.01498 | 0.3991 | 1.0384 | -0.88092 | -0.966 | -0.89702 | -0.93027 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.808 | 0.827 | 0.827 | 0.8 |  |  | 3224 | 3314 | 3325 | 3297 |  |  | 0.498 | 0.517 | 0.517 | 0.49 | --- | --- | 2885.25 | 2975.25 | 2986.25 | 2958.25 | --- | --- | 0.41318 | 0.41041 | 0.41193 | 0.43055 | --- | --- | 0.41652 | 0.0094239 | 0.41644 | 1.0226 | -0.88387 | -0.89059 | -0.8869 | -0.84268 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.821 | 0.836 | 0.841 | 0.813 |  |  | 259 | 252 | 261 | 252 |  |  | 0.511 | 0.526 | 0.531 | 0.503 | --- | --- | -79.75 | -86.75 | -77.75 | -86.75 | --- | --- | -0.01113 | -0.011762 | -0.010442 | -0.0123 | --- | --- | -0.011408 | 0.00080207 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.837 | 0.844 | 0.846 | 0.854 |  |  | 254 | 258 | 257 | 258 |  |  | 0.527 | 0.534 | 0.536 | 0.544 | --- | --- | -84.75 | -80.75 | -81.75 | -80.75 | --- | --- | -0.011469 | -0.010784 | -0.010877 | -0.010586 | --- | --- | -0.010929 | 0.00037977 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.72 | 0.744 | 0.753 | 0.708 |  |  | 2687 | 2715 | 2723 | 2720 |  |  | 0.41 | 0.434 | 0.443 | 0.398 | --- | --- | 2348.25 | 2376.25 | 2384.25 | 2381.25 | --- | --- | 0.40846 | 0.39047 | 0.38383 | 0.42669 | --- | --- | 0.40236 | 0.019268 | 0.40202 | 1.0487 | -0.89536 | -0.9404 | -0.95756 | -0.8517 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.713 | 0.688 | 0.734 | 0.749 |  |  | 2490 | 2549 | 2533 | 2524 |  |  | 0.403 | 0.378 | 0.424 | 0.439 | --- | --- | 2151.25 | 2210.25 | 2194.25 | 2185.25 | --- | --- | 0.38069 | 0.417 | 0.36907 | 0.355 | --- | --- | 0.38044 | 0.026542 | 0.37976 | 1.071 | -0.96576 | -0.87467 | -0.99677 | -1.0356 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.852 | 0.858 | 0.823 | 0.829 |  |  | 557 | 561 | 563 | 592 |  |  | 0.542 | 0.548 | 0.513 | 0.519 | --- | --- | 218.25 | 222.25 | 224.25 | 253.25 | --- | --- | 0.028717 | 0.028923 | 0.031175 | 0.034799 | --- | --- | 0.030904 | 0.0028255 | 0.03081 | 1.0934 | -3.5503 | -3.5431 | -3.4681 | -3.3582 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.841 | 0.824 | 0.807 | 0.862 |  |  | 641 | 658 | 650 | 647 |  |  | 0.531 | 0.514 | 0.497 | 0.552 | --- | --- | 302.25 | 319.25 | 311.25 | 308.25 | --- | --- | 0.040594 | 0.044295 | 0.044662 | 0.039825 | --- | --- | 0.042344 | 0.0024894 | 0.042289 | 1.0607 | -3.2041 | -3.1169 | -3.1086 | -3.2233 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.89 | 0.87 | 0.878 | 0.852 |  |  | 249 | 235 | 235 | 240 |  |  | 0.58 | 0.56 | 0.568 | 0.542 | --- | --- | -89.75 | -103.75 | -103.75 | -98.75 | --- | --- | -0.011036 | -0.013213 | -0.013027 | -0.012993 | --- | --- | -0.012567 | 0.0010255 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.853 | 0.813 | 0.876 | 0.843 |  |  | 229 | 237 | 239 | 234 |  |  | 0.543 | 0.503 | 0.566 | 0.533 | --- | --- | -109.75 | -101.75 | -99.75 | -104.75 | --- | --- | -0.014414 | -0.014426 | -0.012569 | -0.014016 | --- | --- | -0.013856 | 0.00087939 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---