## OD600 reference point
| LUDOX-HS40 | H2O
Replicate 1 | 0.061 | 0.039 | Enter Abs600 absorbance measurements into blue cells
Replicate 2 | 0.062 | 0.039 | Gold cells are calculated
Replicate 3 | 0.063 | 0.043
Replicate 4 | 0.061 | 0.039
Arith. Mean | 0.06175 | 0.04
Corrected Abs600 | 0.02175 |  | Corrected value is particle-only contribution
Reference OD600 | 0.0425 |  | Reference value is for 100uL of LUDOX-HS40 in a well of a standard 96-well flat-bottom plate
OD600/Abs600 | 1.954 |  | Corrected value = scaling factor * measured value

## Fluorescein standard curve
uM Fluorescein | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625 | 0.78125 | 0.39062 | 0.19531 | 0.097656 | 0.048828 | 0
Replicate 1 | 9.167 | 10.099 | 8.136 | 3.694 | 4.01 | 2.294 | 1.263 | 0.983 | 0.373 | 0.592 | 0.222 | -0.023 | Enter fluorescence measurements into blue cells
Replicate 2 | 11.393 | 9.779 | 7.816 | 7.466 | 4.847 | 2.392 | -0.028 | 1.651 | 1.091 | 0.514 | 0.439 | 0.097 | Gold cells are calculated
Replicate 3 | 10.944 | 7.577 | 4.492 | 4.709 | 3.648 | 1.98 | 1.079 | 0.512 | 1.039 | 0.696 | 0.329 | -0.046
Replicate 4 | 8.969 | 8.87 | 5.894 | 5.118 | 4.076 | 2.264 | 1.473 | 1.201 | 0.773 | 0.186 | 0.317 | 0.039 | Values measured are fluorescence from 100uL of X uM fluorescein solution
Arith. Mean | 10.1183 | 9.0812 | 6.5845 | 5.2468 | 4.1452 | 2.2325 | 0.94675 | 1.0867 | 0.819 | 0.497 | 0.32675 | 0.01675
Arith. Std.Dev. | 1.2292 | 1.1299 | 1.7107 | 1.596 | 0.50425 | 0.17698 | 0.66947 | 0.47347 | 0.32834 | 0.22033 | 0.088831 | 0.064428
Values should form a straight line on both linear and log scale
Slope should be 1:1
Common problems:
* Consistent pipetting error --> log graph is a straight line but not 1:1 slope
* Oversaturated detector --> low concentrations linear, but high concentrations saturate or fall
uM Fluorescein/a.u. | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625 | 0.78125 | 0.39062 | 0.19531 | 0.097656 | 0.048828
Mean um Fluorescein/a.u. | 4.9416 | 2.7529 | 1.8984 | 1.1912 | 0.75387 | 0.69989 | 0.82519 | 0.35944 | 0.23848 | 0.19649 | 0.14944
Mean of med-high levels: |  | 1.4593
Final scaling level determined from medium-high points likely to be less impacted by saturation or pipetting error
If needed, you can shift which points are used, but it is likely better to correct instrument settings and protocol.

## Raw Plate Reader Measurements
Raw Plate Readings |  | If you followed the recommended plate layout:
Copy fluorescence and Abs600 measurements from your plate reader into blue cells
They will automatically propagate into the correct locations in the Fluorescence Measurement Sheet
Fluorescence Raw Readings: | [9 empty] | Abs600 Raw Readings:
Hour 0: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 0: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | -0.017 | 0.089 | 0.311 | 0.205 | 0.055 | 0.156 | -0.057 | 0.119 | 0.006 | Colony 1, Replicate 1 | 0.025 | 0.046 | 0.044 | 0.046 | 0.052 | 0.052 | 0.054 | 0.056 | 0.089
Colony 1, Replicate 2 | 0.09 | 0.006 | 0.174 | 0.175 | 0.161 | 0 | 0.022 | 0.136 | 0.125 | Colony 1, Replicate 2 | -0.073 | 0.049 | 0.047 | 0.05 | 0.054 | 0.059 | 0.059 | 0.06 | 0.062
Colony 1, Replicate 3 | 0.145 | 0.006 | 0.017 | -0.045 | 0.074 | 0.051 | 0.006 | -0.022 | 0.068 | Colony 1, Replicate 3 | 0.082 | 0.047 | 0.046 | 0.047 | 0.05 | 0.05 | 0.051 | 0.05 | 0.047
Colony 1, Replicate 4 | 0.006 | 0.067 | -0.012 | 0.05 | 0.051 | 0.033 | 0.05 | 0.011 | 0.034 | Colony 1, Replicate 4 | 0.025 | 0.053 | 0.046 | 0.044 | 0.049 | 0.048 | 0.053 | 0.052 | 0.05
Colony 2, Replicate 1 | 0.011 | -0.011 | 0.017 | 0.051 | 0.085 | 0.057 | -0.045 | -0.011 | -0.023 | Colony 2, Replicate 1 | 0.034 | 0.047 | 0.047 | 0.048 | 0.053 | 0.05 | 0.051 | 0.05 | 0.046
Colony 2, Replicate 2 | 0.034 | 0.079 | 0.085 | 0.183 | -0.029 | -0.143 | -0.034 | 0.006 | 0.146 | Colony 2, Replicate 2 | 0.052 | 0.048 | 0.048 | 0.048 | 0.075 | 0.05 | 0.053 | 0.05 | 0.052
Colony 2, Replicate 3 | 0.011 | 0.029 | -0.079 | 0.034 | 0.045 | -0.045 | 0.09 | 0.023 | 0.051 | Colony 2, Replicate 3 | 0.048 | 0.048 | 0.049 | 0.049 | 0.055 | 0.048 | 0.051 | 0.05 | 0.048
Colony 2, Replicate 4 | 0.034 | 0.346 | 0.022 | 0.106 | 0.033 | -0.04 | 0.073 | -0.011 | -0.006 | Colony 2, Replicate 4 | 0.046 | 0.047 | 0.051 | 0.048 | 0.05 | 0.046 | 0.048 | 0.044 | 0.045
Hour 2: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 2: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 0 | 0.04 | 0.165 | 0.179 | 0.04 | 0.292 | 0.085 | 0 | 0.023 | Colony 1, Replicate 1 | 0.068 | 0.069 | 0.053 | 0.061 | 0.072 | 0.059 | 0.073 | 0.077 | 0.056
Colony 1, Replicate 2 | 0.119 | 0.147 | 0.268 | 0.276 | 0.006 | 0.063 | 0.057 | 0.062 | 0.097 | Colony 1, Replicate 2 | 0.066 | 0.071 | 0.053 | 0.061 | 0.072 | 0.057 | 0.068 | 0.077 | 0.055
Colony 1, Replicate 3 | 0.218 | 0.04 | -0.023 | 0.046 | 0.034 | 0.035 | -0.028 | 0.09 | 0.067 | Colony 1, Replicate 3 | 0.068 | 0.072 | 0.053 | 0.059 | 0.077 | 0.056 | 0.072 | 0.077 | 0.05
Colony 1, Replicate 4 | 0.057 | 0.133 | 0.023 | 0.154 | 0.062 | 0.029 | 0.136 | -0.011 | 0.023 | Colony 1, Replicate 4 | 0.069 | 0.07 | 0.053 | 0.059 | 0.073 | 0.054 | 0.071 | 0.074 | 0.051
Colony 2, Replicate 1 | 0.116 | 0.04 | 0.017 | 0.529 | 0.169 | 0.197 | 0.045 | 0.017 | -0.111 | Colony 2, Replicate 1 | 0.063 | 0.069 | 0.053 | 0.071 | 0.077 | 0.057 | 0.059 | 0.065 | 0.043
Colony 2, Replicate 2 | 0.148 | 0.103 | 0.057 | 0.317 | 0.063 | -0.006 | 0.186 | -0.017 | 0.063 | Colony 2, Replicate 2 | 0.066 | 0.072 | 0.056 | 0.073 | 0.079 | 0.058 | 0.067 | 0.071 | 0.047
Colony 2, Replicate 3 | -0.011 | -0.069 | -0.056 | 0.218 | 0.199 | 0.084 | 0.468 | 0.011 | 0.284 | Colony 2, Replicate 3 | 0.069 | 0.075 | 0.057 | 0.075 | 0.079 | 0.057 | 0.068 | 0.071 | 0.043
Colony 2, Replicate 4 | -0.017 | 0.04 | 0.056 | 0.193 | 0.09 | 0.149 | 0.063 | 0.122 | 0.092 | Colony 2, Replicate 4 | 0.065 | 0.069 | 0.058 | 0.073 | 0.074 | 0.054 | 0.063 | 0.066 | 0.044
Hour 4: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 4: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 0.18 | 0.138 | 0.011 | 0.125 | -0.017 | 0.124 | -0.034 | 0.091 | 0.053 | Colony 1, Replicate 1 | 0.094 | 0.126 | 0.059 | 0.112 | 0.104 | 0.077 | 0.101 | 0.126 | 0.057
Colony 1, Replicate 2 | 0.367 | 0.08 | 0.028 | 0.157 | 0.115 | 0.085 | 0.135 | 0.046 | 0.194 | Colony 1, Replicate 2 | 0.102 | 0.135 | 0.063 | 0.109 | 0.11 | 0.08 | 0.117 | 0.114 | 0.057
Colony 1, Replicate 3 | 0.05 | 0.017 | 0.017 | 0.138 | 0.118 | 0.034 | 0.216 | 0.045 | 0.035 | Colony 1, Replicate 3 | 0.096 | 0.122 | 0.056 | 0.1 | 0.105 | 0.066 | 0.091 | 0.121 | 0.051
Colony 1, Replicate 4 | 0.152 | 0.126 | 0.011 | 0.151 | 0.121 | 0.144 | 0.283 | 0.023 | 0.103 | Colony 1, Replicate 4 | 0.101 | 0.132 | 0.058 | 0.105 | 0.114 | 0.071 | 0.118 | 0.132 | 0.052
Colony 2, Replicate 1 | 0.04 | 0.079 | 0.119 | 0.507 | -0.023 | 0.411 | 0.069 | 0.035 | 0.053 | Colony 2, Replicate 1 | 0.07 | 0.133 | 0.056 | 0.129 | 0.117 | 0.079 | 0.107 | 0.129 | 0.057
Colony 2, Replicate 2 | 0.192 | 0.034 | 0.029 | 0.337 | 0.006 | 0.541 | 0.063 | 0.006 | 0.194 | Colony 2, Replicate 2 | 0.068 | 0.147 | 0.056 | 0.125 | 0.13 | 0.082 | 0.11 | 0.137 | 0.057
Colony 2, Replicate 3 | 0.017 | -0.045 | -0.028 | 0.296 | 0.145 | 0.381 | 0.118 | -0.133 | 0.035 | Colony 2, Replicate 3 | 0.075 | 0.156 | 0.061 | 0.137 | 0.142 | 0.081 | 0.117 | 0.134 | 0.051
Colony 2, Replicate 4 | 0.327 | 0.052 | 0.102 | 0.325 | 0.045 | 0.097 | 0.349 | -0.023 | 0.103 | Colony 2, Replicate 4 | 0.069 | 0.149 | 0.062 | 0.13 | 0.133 | 0.079 | 0.114 | 0.126 | 0.052
Hour 6: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 6: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 0.141 | 0.045 | 0.028 | 0.302 | 0.057 | 0.224 | 0.15 | -0.108 | 0.165 | Colony 1, Replicate 1 | 0.143 | 0.141 | 0.054 | 0.112 | 0.11 | 0.086 | 0.113 | 0.128 | 0.062
Colony 1, Replicate 2 | 0 | 0.046 | 0.028 | 0.441 | -0.029 | 0.023 | 0.091 | -0.051 | 0.08 | Colony 1, Replicate 2 | 0.149 | 0.153 | 0.058 | 0.112 | 0.12 | 0.098 | 0.108 | 0.117 | 0.066
Colony 1, Replicate 3 | 0.39 | 0.115 | -0.104 | 0.012 | 0.075 | 0.122 | 0.006 | 0.096 | 0.006 | Colony 1, Replicate 3 | 0.136 | 0.139 | 0.052 | 0.12 | 0.121 | 0.088 | 0.109 | 0.125 | 0.05
Colony 1, Replicate 4 | 0.547 | 0.145 | -0.023 | 0.306 | 0.121 | 0.098 | 0.206 | -0.058 | 0.052 | Colony 1, Replicate 4 | 0.161 | 0.184 | 0.054 | 0.13 | 0.125 | 0.1 | 0.124 | 0.13 | 0.053
Colony 2, Replicate 1 | 0.08 | 0.093 | -0.011 | 0.681 | 0.018 | 0.149 | 0.171 | 0.006 | 0.219 | Colony 2, Replicate 1 | 0.066 | 0.152 | 0.054 | 0.128 | 0.127 | 0.089 | 0.123 | 0.13 | 0.051
Colony 2, Replicate 2 | 0.144 | 0.12 | 0.08 | 0.592 | -0.035 | 0.433 | 0.086 | -0.041 | 0.051 | Colony 2, Replicate 2 | 0.065 | 0.15 | 0.054 | 0.142 | 0.136 | 0.09 | 0.121 | 0.124 | 0.056
Colony 2, Replicate 3 | 0.006 | 0.058 | 0.046 | 0.4 | -0.085 | 0.202 | -0.017 | 0.122 | 0.023 | Colony 2, Replicate 3 | 0.077 | 0.17 | 0.058 | 0.136 | 0.14 | 0.093 | 0.123 | 0.132 | 0.057
Colony 2, Replicate 4 | 0.125 | 0.069 | -0.069 | 0.396 | 0.058 | 0.472 | 0.084 | -0.227 | -0.116 | Colony 2, Replicate 4 | 0.067 | 0.19 | 0.057 | 0.153 | 0.143 | 0.097 | 0.136 | 0.124 | 0.048
Plate pattern:
| A1 | A2 | A3 | A4 | A5 | A6 | A7 | A8 | A9
| B1 | B2 | B3 | B4 | B5 | B6 | B7 | B8 | B9
| C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9
| D1 | D2 | D3 | D4 | D5 | D6 | D7 | D8 | D9
| E1 | E2 | E3 | E4 | E5 | E6 | E7 | E8 | E9
| F1 | F2 | F3 | F4 | F5 | F6 | F7 | F8 | F9
| G1 | G2 | G3 | G4 | G5 | G6 | G7 | G8 | G9
| H1 | H2 | H3 | H4 | H5 | H6 | H7 | H8 | H9

## Fluorescence Measurement
Unit Scaling Factors: | These are imported from the prior two sheets | [5 empty] | Enter fluorescence and Abs600 measurements into blue cells on "Raw Plate Reader Measurements"
OD600/Abs600 | 1.954 | [5 empty] | They will be copied into the green cells on this sheet.
uM Fluorescein/a.u. | 1.4593 | [5 empty] | Gold cells are calculated
If you have more replicates, unhide the extra columns
Experimental Values: | Raw Abs600 | [5 empty] | Raw Fluorescence
Sample set: | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6
Blank media | 0.089 | 0.062 | 0.047 | 0.05 |  |  | 0.006 | 0.125 | 0.068 | 0.034
Blank mean: | 0.062 | [5 empty] | 0.05825 | [5 empty] | OD - Background | [5 empty] | Fluorescence - Background | [5 empty] | uM Fluorescein / OD600 | [5 empty] | Summary Statistics |  |  |  | Ln uM FITC / OD600
Hour 0: | [12 empty] | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Arith. Mean | Arith. Std.Dev. | Geo. Mean | Geo. Std. Dev. | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6
Negative Control (Colony 1) | 0.025 | -0.073 | 0.082 | 0.025 |  |  | -0.017 | 0.09 | 0.145 | 0.006 |  |  | -0.037 | -0.135 | 0.02 | -0.037 | --- | --- | -0.07525 | 0.03175 | 0.08675 | -0.05225 | --- | --- | 1.5188 | -0.17564 | 3.2392 | 1.0546 | --- | --- | 1.4093 | 1.414 | #NUM! | #NUM! | 0.41794 | #NUM! | 1.1753 | 0.053161 | --- | ---
Negative Control (Colony 2) | 0.034 | 0.052 | 0.048 | 0.046 |  |  | 0.011 | 0.034 | 0.011 | 0.034 |  |  | -0.028 | -0.01 | -0.014 | -0.016 | --- | --- | -0.04725 | -0.02425 | -0.04725 | -0.02425 | --- | --- | 1.2602 | 1.811 | 2.5204 | 1.1319 | --- | --- | 1.6809 | 0.6325 | 1.5974 | 1.4398 | 0.23129 | 0.59387 | 0.92443 | 0.12387 | --- | ---
Positive Control (Colony 1) | 0.046 | 0.049 | 0.047 | 0.053 |  |  | 0.089 | 0.006 | 0.006 | 0.067 |  |  | -0.016 | -0.013 | -0.015 | -0.009 | --- | --- | 0.03075 | -0.05225 | -0.05225 | 0.00875 | --- | --- | -1.4353 | 3.0016 | 2.6013 | -0.72605 | --- | --- | 0.8604 | 2.2659 | #NUM! | #NUM! | #NUM! | 1.0991 | 0.95603 | #NUM! | --- | ---
Positive Control (Colony 2) | 0.047 | 0.048 | 0.048 | 0.047 |  |  | -0.011 | 0.079 | 0.029 | 0.346 |  |  | -0.015 | -0.014 | -0.014 | -0.015 | --- | --- | -0.06925 | 0.02075 | -0.02925 | 0.28775 | --- | --- | 3.4477 | -1.1069 | 1.5603 | -14.3261 | --- | --- | -2.6062 | 8.0335 | #NUM! | #NUM! | 1.2377 | #NUM! | 0.44486 | #NUM! | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.044 | 0.047 | 0.046 | 0.046 |  |  | 0.311 | 0.174 | 0.017 | -0.012 |  |  | -0.018 | -0.015 | -0.016 | -0.016 | --- | --- | 0.25275 | 0.11575 | -0.04125 | -0.07025 | --- | --- | -10.4863 | -5.7628 | 1.9253 | 3.2789 | --- | --- | -2.7612 | 6.5098 | #NUM! | #NUM! | #NUM! | #NUM! | 0.6551 | 1.1875 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.047 | 0.048 | 0.049 | 0.051 |  |  | 0.017 | 0.085 | -0.079 | 0.022 |  |  | -0.015 | -0.014 | -0.013 | -0.011 | --- | --- | -0.04125 | 0.02675 | -0.13725 | -0.03625 | --- | --- | 2.0537 | -1.4269 | 7.8845 | 2.461 | --- | --- | 2.7431 | 3.8461 | #NUM! | #NUM! | 0.71964 | #NUM! | 2.0649 | 0.90058 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.046 | 0.05 | 0.047 | 0.044 |  |  | 0.205 | 0.175 | -0.045 | 0.05 |  |  | -0.016 | -0.012 | -0.015 | -0.018 | --- | --- | 0.14675 | 0.11675 | -0.10325 | -0.00825 | --- | --- | -6.8495 | -7.2657 | 5.1405 | 0.34228 | --- | --- | -2.1581 | 5.9894 | #NUM! | #NUM! | #NUM! | #NUM! | 1.6371 | -1.0721 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.048 | 0.048 | 0.049 | 0.048 |  |  | 0.051 | 0.183 | 0.034 | 0.106 |  |  | -0.014 | -0.014 | -0.013 | -0.014 | --- | --- | -0.00725 | 0.12475 | -0.02425 | 0.04775 | --- | --- | 0.38673 | -6.6545 | 1.3931 | -2.5471 | --- | --- | -1.8555 | 3.6097 | #NUM! | #NUM! | -0.95002 | #NUM! | 0.33151 | #NUM! | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.052 | 0.054 | 0.05 | 0.049 |  |  | 0.055 | 0.161 | 0.074 | 0.051 |  |  | -0.01 | -0.008 | -0.012 | -0.013 | --- | --- | -0.00325 | 0.10275 | 0.01575 | -0.00725 | --- | --- | 0.24271 | -9.5917 | -0.98017 | 0.41648 | --- | --- | -2.4782 | 4.7829 | #NUM! | #NUM! | -1.4159 | #NUM! | #NUM! | -0.87591 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.053 | 0.075 | 0.055 | 0.05 |  |  | 0.085 | -0.029 | 0.045 | 0.033 |  |  | -0.009 | 0.013 | -0.007 | -0.012 | --- | --- | 0.02675 | -0.08725 | -0.01325 | -0.02525 | --- | --- | -2.2196 | -5.0122 | 1.4136 | 1.5714 | --- | --- | -1.0617 | 3.1627 | #NUM! | #NUM! | #NUM! | #NUM! | 0.34613 | 0.45196 | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.052 | 0.059 | 0.05 | 0.048 |  |  | 0.156 | 0 | 0.051 | 0.033 |  |  | -0.01 | -0.003 | -0.012 | -0.014 | --- | --- | 0.09775 | -0.05825 | -0.00725 | -0.02525 | --- | --- | -7.2999 | 14.5003 | 0.45119 | 1.3469 | --- | --- | 2.2496 | 9.0429 | #NUM! | #NUM! | #NUM! | 2.6742 | -0.79587 | 0.29781 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.05 | 0.05 | 0.048 | 0.046 |  |  | 0.057 | -0.143 | -0.045 | -0.04 |  |  | -0.012 | -0.012 | -0.014 | -0.016 | --- | --- | -0.00125 | -0.20125 | -0.10325 | -0.09825 | --- | --- | 0.077791 | 12.5244 | 5.5076 | 4.5858 | --- | --- | 5.6739 | 5.1464 | 2.2272 | 9.763 | -2.5537 | 2.5277 | 1.7061 | 1.523 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.054 | 0.059 | 0.051 | 0.053 |  |  | -0.057 | 0.022 | 0.006 | 0.05 |  |  | -0.008 | -0.003 | -0.011 | -0.009 | --- | --- | -0.11525 | -0.03625 | -0.05225 | -0.00825 | --- | --- | 10.7586 | 9.0238 | 3.5473 | 0.68456 | --- | --- | 6.0036 | 4.6924 | 3.9184 | 3.5286 | 2.3757 | 2.1999 | 1.2662 | -0.37897 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.051 | 0.053 | 0.051 | 0.048 |  |  | -0.045 | -0.034 | 0.09 | 0.073 |  |  | -0.011 | -0.009 | -0.011 | -0.014 | --- | --- | -0.10325 | -0.09225 | 0.03175 | 0.01475 | --- | --- | 7.0097 | 7.6547 | -2.1555 | -0.7868 | --- | --- | 2.9305 | 5.12 | #NUM! | #NUM! | 1.9473 | 2.0353 | #NUM! | #NUM! | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.056 | 0.06 | 0.05 | 0.052 |  |  | 0.119 | 0.136 | -0.022 | 0.011 |  |  | -0.006 | -0.002 | -0.012 | -0.01 | --- | --- | 0.06075 | 0.07775 | -0.08025 | -0.04725 | --- | --- | -7.5613 | -29.0318 | 4.9942 | 3.5286 | --- | --- | -7.0176 | 15.7101 | #NUM! | #NUM! | #NUM! | #NUM! | 1.6083 | 1.2609 | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.05 | 0.05 | 0.05 | 0.044 |  |  | -0.011 | 0.006 | 0.023 | -0.011 |  |  | -0.012 | -0.012 | -0.012 | -0.018 | --- | --- | -0.06925 | -0.05225 | -0.03525 | -0.06925 | --- | --- | 4.3096 | 3.2517 | 2.1937 | 2.8731 | --- | --- | 3.157 | 0.88432 | 3.0656 | 1.3235 | 1.4609 | 1.1792 | 0.7856 | 1.0554 | --- | ---
Hour 2:
Negative Control (Colony 1) | 0.068 | 0.066 | 0.068 | 0.069 |  |  | 0 | 0.119 | 0.218 | 0.057 |  |  | 0.006 | 0.004 | 0.006 | 0.007 | --- | --- | -0.05825 | 0.06075 | 0.15975 | -0.00125 | --- | --- | -7.2502 | 11.342 | 19.8835 | -0.13336 | --- | --- | 5.9605 | 12.0342 | #NUM! | #NUM! | #NUM! | 2.4285 | 2.9899 | #NUM! | --- | ---
Negative Control (Colony 2) | 0.063 | 0.066 | 0.069 | 0.065 |  |  | 0.116 | 0.148 | -0.011 | -0.017 |  |  | 0.001 | 0.004 | 0.007 | 0.003 | --- | --- | 0.05775 | 0.08975 | -0.06925 | -0.07525 | --- | --- | 43.1276 | 16.7563 | -7.388 | -18.7322 | --- | --- | 8.4409 | 27.4544 | #NUM! | #NUM! | 3.7642 | 2.8188 | #NUM! | #NUM! | --- | ---
Positive Control (Colony 1) | 0.069 | 0.071 | 0.072 | 0.07 |  |  | 0.04 | 0.147 | 0.04 | 0.133 |  |  | 0.007 | 0.009 | 0.01 | 0.008 | --- | --- | -0.01825 | 0.08875 | -0.01825 | 0.07475 | --- | --- | -1.947 | 7.3643 | -1.3629 | 6.9779 | --- | --- | 2.7581 | 5.1037 | #NUM! | #NUM! | #NUM! | 1.9966 | #NUM! | 1.9427 | --- | ---
Positive Control (Colony 2) | 0.069 | 0.072 | 0.075 | 0.069 |  |  | 0.04 | 0.103 | -0.069 | 0.04 |  |  | 0.007 | 0.01 | 0.013 | 0.007 | --- | --- | -0.01825 | 0.04475 | -0.12725 | -0.01825 | --- | --- | -1.947 | 3.3419 | -7.31 | -1.947 | --- | --- | -1.9655 | 4.3487 | #NUM! | #NUM! | #NUM! | 1.2065 | #NUM! | #NUM! | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.053 | 0.053 | 0.053 | 0.053 |  |  | 0.165 | 0.268 | -0.023 | 0.023 |  |  | -0.009 | -0.009 | -0.009 | -0.009 | --- | --- | 0.10675 | 0.20975 | -0.08125 | -0.03525 | --- | --- | -8.8579 | -17.4045 | 6.7419 | 2.925 | --- | --- | -4.1489 | 11.0534 | #NUM! | #NUM! | #NUM! | #NUM! | 1.9083 | 1.0733 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.053 | 0.056 | 0.057 | 0.058 |  |  | 0.017 | 0.057 | -0.056 | 0.056 |  |  | -0.009 | -0.006 | -0.005 | -0.004 | --- | --- | -0.04125 | -0.00125 | -0.11425 | -0.00225 | --- | --- | 3.4228 | 0.15558 | 17.0643 | 0.42007 | --- | --- | 5.2657 | 8.0041 | 1.3978 | 8.2298 | 1.2305 | -1.8606 | 2.837 | -0.86733 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.061 | 0.061 | 0.059 | 0.059 |  |  | 0.179 | 0.276 | 0.046 | 0.154 |  |  | -0.001 | -0.001 | -0.003 | -0.003 | --- | --- | 0.12075 | 0.21775 | -0.01225 | 0.09575 | --- | --- | -90.1758 | -162.6152 | 3.0494 | -23.8353 | --- | --- | -68.3942 | 74.0309 | #NUM! | #NUM! | #NUM! | #NUM! | 1.115 | #NUM! | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.071 | 0.073 | 0.075 | 0.073 |  |  | 0.529 | 0.317 | 0.218 | 0.193 |  |  | 0.009 | 0.011 | 0.013 | 0.011 | --- | --- | 0.47075 | 0.25875 | 0.15975 | 0.13475 | --- | --- | 39.0617 | 17.5667 | 9.177 | 9.1483 | --- | --- | 18.7384 | 14.1162 | 15.4925 | 1.991 | 3.6651 | 2.866 | 2.2167 | 2.2136 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.072 | 0.072 | 0.077 | 0.073 |  |  | 0.04 | 0.006 | 0.034 | 0.062 |  |  | 0.01 | 0.01 | 0.015 | 0.011 | --- | --- | -0.01825 | -0.05225 | -0.02425 | 0.00375 | --- | --- | -1.3629 | -3.902 | -1.2073 | 0.25459 | --- | --- | -1.5544 | 1.7264 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | -1.3681 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.077 | 0.079 | 0.079 | 0.074 |  |  | 0.169 | 0.063 | 0.199 | 0.09 |  |  | 0.015 | 0.017 | 0.017 | 0.012 | --- | --- | 0.11075 | 0.00475 | 0.14075 | 0.03175 | --- | --- | 5.5139 | 0.20866 | 6.183 | 1.9759 | --- | --- | 3.4704 | 2.8523 | 1.9363 | 4.8128 | 1.7073 | -1.567 | 1.8218 | 0.68103 | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.059 | 0.057 | 0.056 | 0.054 |  |  | 0.292 | 0.063 | 0.035 | 0.029 |  |  | -0.003 | -0.005 | -0.006 | -0.008 | --- | --- | 0.23375 | 0.00475 | -0.02325 | -0.02925 | --- | --- | -58.188 | -0.70946 | 2.8938 | 2.7305 | --- | --- | -13.3183 | 29.9592 | #NUM! | #NUM! | #NUM! | #NUM! | 1.0626 | 1.0045 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.057 | 0.058 | 0.057 | 0.054 |  |  | 0.197 | -0.006 | 0.084 | 0.149 |  |  | -0.005 | -0.004 | -0.005 | -0.008 | --- | --- | 0.13875 | -0.06425 | 0.02575 | 0.09075 | --- | --- | -20.7236 | 11.9954 | -3.846 | -8.4715 | --- | --- | -5.2614 | 13.5301 | #NUM! | #NUM! | #NUM! | 2.4845 | #NUM! | #NUM! | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.073 | 0.068 | 0.072 | 0.071 |  |  | 0.085 | 0.057 | -0.028 | 0.136 |  |  | 0.011 | 0.006 | 0.01 | 0.009 | --- | --- | 0.02675 | -0.00125 | -0.08625 | 0.07775 | --- | --- | 1.8161 | -0.15558 | -6.4411 | 6.4515 | --- | --- | 0.41772 | 5.3459 | #NUM! | #NUM! | 0.59668 | #NUM! | #NUM! | 1.8643 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.059 | 0.067 | 0.068 | 0.063 |  |  | 0.045 | 0.186 | 0.468 | 0.063 |  |  | -0.003 | 0.005 | 0.006 | 0.001 | --- | --- | -0.01325 | 0.12775 | 0.40975 | 0.00475 | --- | --- | 3.2984 | 19.0807 | 51.0001 | 3.5473 | --- | --- | 19.2316 | 22.4286 | 10.3297 | 3.812 | 1.1934 | 2.9487 | 3.9318 | 1.2662 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.077 | 0.077 | 0.077 | 0.074 |  |  | 0 | 0.062 | 0.09 | -0.011 |  |  | 0.015 | 0.015 | 0.015 | 0.012 | --- | --- | -0.05825 | 0.00375 | 0.03175 | -0.06925 | --- | --- | -2.9001 | 0.1867 | 1.5807 | -4.3096 | --- | --- | -1.3606 | 2.7149 | #NUM! | #NUM! | #NUM! | -1.6783 | 0.45788 | #NUM! | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.065 | 0.071 | 0.071 | 0.066 |  |  | 0.017 | -0.017 | 0.011 | 0.122 |  |  | 0.003 | 0.009 | 0.009 | 0.004 | --- | --- | -0.04125 | -0.07525 | -0.04725 | 0.06375 | --- | --- | -10.2685 | -6.2441 | -3.9207 | 11.9021 | --- | --- | -2.1328 | 9.7171 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | 2.4767 | --- | ---
Hour 4:
Negative Control (Colony 1) | 0.094 | 0.102 | 0.096 | 0.101 |  |  | 0.18 | 0.367 | 0.05 | 0.152 |  |  | 0.032 | 0.04 | 0.034 | 0.039 | --- | --- | 0.12175 | 0.30875 | -0.00825 | 0.09375 | --- | --- | 2.8413 | 5.7643 | -0.18121 | #NUM! | --- | --- | #NUM! | #NUM! | #NUM! | #NUM! | 1.0443 | 1.7517 | #NUM! | --- | --- | ---
Negative Control (Colony 2) | 0.07 | 0.068 | 0.075 | 0.069 |  |  | 0.04 | 0.192 | 0.017 | 0.327 |  |  | 0.008 | 0.006 | 0.013 | 0.007 | --- | --- | -0.01825 | 0.13375 | -0.04125 | 0.26875 | --- | --- | -1.7036 | 16.6474 | -2.3696 | #NUM! | --- | --- | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | 2.8123 | #NUM! | --- | --- | ---
Positive Control (Colony 1) | 0.126 | 0.135 | 0.122 | 0.132 |  |  | 0.138 | 0.08 | 0.017 | 0.126 |  |  | 0.064 | 0.073 | 0.06 | 0.07 | --- | --- | 0.07975 | 0.02175 | -0.04125 | 0.06775 | --- | --- | 0.93058 | 0.2225 | -0.51342 | #NUM! | --- | --- | #NUM! | #NUM! | #NUM! | #NUM! | -0.071947 | -1.5028 | #NUM! | --- | --- | ---
Positive Control (Colony 2) | 0.133 | 0.147 | 0.156 | 0.149 |  |  | 0.079 | 0.034 | -0.045 | 0.052 |  |  | 0.071 | 0.085 | 0.094 | 0.087 | --- | --- | 0.02075 | -0.02425 | -0.10325 | -0.00625 | --- | --- | 0.21825 | -0.21306 | -0.82029 | #NUM! | --- | --- | #NUM! | #NUM! | #NUM! | #NUM! | -1.5221 | #NUM! | #NUM! | --- | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.059 | 0.063 | 0.056 | 0.058 |  |  | 0.011 | 0.028 | 0.017 | 0.011 |  |  | -0.003 | 0.001 | -0.006 | -0.004 | --- | --- | -0.04725 | -0.03025 | -0.04125 | -0.04725 | --- | --- | 11.7621 | -22.5906 | 5.1342 | #NUM! | --- | --- | #NUM! | #NUM! | #NUM! | #NUM! | 2.4649 | #NUM! | 1.6359 | --- | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.056 | 0.056 | 0.061 | 0.062 |  |  | 0.119 | 0.029 | -0.028 | 0.102 |  |  | -0.006 | -0.006 | -0.001 | 0 | --- | --- | 0.06075 | -0.02925 | -0.08625 | 0.04375 | --- | --- | -7.5613 | 3.6406 | 64.4113 | #DIV/0! | --- | --- | #DIV/0! | #DIV/0! | #DIV/0! | #NUM! | #NUM! | 1.2922 | 4.1653 | --- | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.112 | 0.109 | 0.1 | 0.105 |  |  | 0.125 | 0.157 | 0.138 | 0.151 |  |  | 0.05 | 0.047 | 0.038 | 0.043 | --- | --- | 0.06675 | 0.09875 | 0.07975 | 0.09275 | --- | --- | 0.99697 | 1.5691 | 1.5673 | 1.6108 | --- | --- | 1.436 | 0.2934 | 1.4097 | 1.2602 | -0.0030297 | 0.45048 | 0.44935 | 0.47675 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.129 | 0.125 | 0.137 | 0.13 |  |  | 0.507 | 0.337 | 0.296 | 0.325 |  |  | 0.067 | 0.063 | 0.075 | 0.068 | --- | --- | 0.44875 | 0.27875 | 0.23775 | 0.26675 | --- | --- | 5.0019 | 3.3043 | 2.3673 | 2.9295 | --- | --- | 3.4008 | 1.1347 | 3.272 | 1.3699 | 1.6098 | 1.1952 | 0.86177 | 1.0748 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.104 | 0.11 | 0.105 | 0.114 |  |  | -0.017 | 0.115 | 0.118 | 0.121 |  |  | 0.042 | 0.048 | 0.043 | 0.052 | --- | --- | -0.07525 | 0.05675 | 0.05975 | 0.06275 | --- | --- | -1.338 | 0.88293 | 1.0377 | 0.90118 | --- | --- | 0.37095 | 1.1414 | #NUM! | #NUM! | #NUM! | -0.12451 | 0.037008 | -0.10405 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.117 | 0.13 | 0.142 | 0.133 |  |  | -0.023 | 0.006 | 0.145 | 0.045 |  |  | 0.055 | 0.068 | 0.08 | 0.071 | --- | --- | -0.08125 | -0.05225 | 0.08675 | -0.01325 | --- | --- | -1.1032 | -0.57383 | 0.80981 | -0.13937 | --- | --- | -0.25165 | 0.81 | #NUM! | #NUM! | #NUM! | #NUM! | -0.21096 | #NUM! | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.077 | 0.08 | 0.066 | 0.071 |  |  | 0.124 | 0.085 | 0.034 | 0.144 |  |  | 0.015 | 0.018 | 0.004 | 0.009 | --- | --- | 0.06575 | 0.02675 | -0.02425 | 0.08575 | --- | --- | 3.2735 | 1.1098 | -4.5275 | 7.1153 | --- | --- | 1.7428 | 4.8622 | #NUM! | #NUM! | 1.1858 | 0.1042 | #NUM! | 1.9623 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.079 | 0.082 | 0.081 | 0.079 |  |  | 0.411 | 0.541 | 0.381 | 0.097 |  |  | 0.017 | 0.02 | 0.019 | 0.017 | --- | --- | 0.35275 | 0.48275 | 0.32275 | 0.03875 | --- | --- | 15.4961 | 18.0258 | 12.6857 | 1.7023 | --- | --- | 11.9775 | 7.189 | 8.8128 | 3.0209 | 2.7406 | 2.8918 | 2.5405 | 0.53196 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.101 | 0.117 | 0.091 | 0.118 |  |  | -0.034 | 0.135 | 0.216 | 0.283 |  |  | 0.039 | 0.055 | 0.029 | 0.056 | --- | --- | -0.09225 | 0.07675 | 0.15775 | 0.22475 | --- | --- | -1.7665 | 1.0421 | 4.0623 | 2.9972 | --- | --- | 1.5838 | 2.5598 | #NUM! | #NUM! | #NUM! | 0.041259 | 1.4018 | 1.0977 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.107 | 0.11 | 0.117 | 0.114 |  |  | 0.069 | 0.063 | 0.118 | 0.349 |  |  | 0.045 | 0.048 | 0.055 | 0.052 | --- | --- | 0.01075 | 0.00475 | 0.05975 | 0.29075 | --- | --- | 0.1784 | 0.073902 | 0.81129 | 4.1756 | --- | --- | 1.3098 | 1.9381 | 0.45971 | 5.8871 | -1.7237 | -2.605 | -0.20912 | 1.4293 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.126 | 0.114 | 0.121 | 0.132 |  |  | 0.091 | 0.046 | 0.045 | 0.023 |  |  | 0.064 | 0.052 | 0.059 | 0.07 | --- | --- | 0.03275 | -0.01225 | -0.01325 | -0.03525 | --- | --- | 0.38215 | -0.17593 | -0.16771 | -0.37607 | --- | --- | -0.084389 | 0.32561 | #NUM! | #NUM! | -0.96194 | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.129 | 0.137 | 0.134 | 0.126 |  |  | 0.035 | 0.006 | -0.133 | -0.023 |  |  | 0.067 | 0.075 | 0.072 | 0.064 | --- | --- | -0.02325 | -0.05225 | -0.19125 | -0.08125 | --- | --- | -0.25915 | -0.52027 | -1.9837 | -0.94808 | --- | --- | -0.9278 | 0.75905 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Hour 6:
Negative Control (Colony 1) | 0.143 | 0.149 | 0.136 | 0.161 |  |  | 0.141 | 0 | 0.39 | 0.547 |  |  | 0.081 | 0.087 | 0.074 | 0.099 | --- | --- | 0.08275 | -0.05825 | 0.33175 | 0.48875 | --- | --- | 0.76293 | -0.50001 | 3.348 | 3.6868 | --- | --- | 1.8244 | 2.0265 | #NUM! | #NUM! | -0.27059 | #NUM! | 1.2084 | 1.3048 | --- | ---
Negative Control (Colony 2) | 0.066 | 0.065 | 0.077 | 0.067 |  |  | 0.08 | 0.144 | 0.006 | 0.125 |  |  | 0.004 | 0.003 | 0.015 | 0.005 | --- | --- | 0.02175 | 0.08575 | -0.05225 | 0.06675 | --- | --- | 4.0607 | 21.346 | -2.6013 | 9.9697 | --- | --- | 8.1938 | 10.1612 | #NUM! | #NUM! | 1.4014 | 3.0609 | #NUM! | 2.2996 | --- | ---
Positive Control (Colony 1) | 0.141 | 0.153 | 0.139 | 0.184 |  |  | 0.045 | 0.046 | 0.115 | 0.145 |  |  | 0.079 | 0.091 | 0.077 | 0.122 | --- | --- | -0.01325 | -0.01225 | 0.05675 | 0.08675 | --- | --- | -0.12525 | -0.10053 | 0.5504 | 0.53102 | --- | --- | 0.21391 | 0.37758 | #NUM! | #NUM! | #NUM! | #NUM! | -0.59711 | -0.63295 | --- | ---
Positive Control (Colony 2) | 0.152 | 0.15 | 0.17 | 0.19 |  |  | 0.093 | 0.12 | 0.058 | 0.069 |  |  | 0.09 | 0.088 | 0.108 | 0.128 | --- | --- | 0.03475 | 0.06175 | -0.00025 | 0.01075 | --- | --- | 0.28835 | 0.52403 | -0.0017287 | 0.062719 | --- | --- | 0.21834 | 0.23874 | #NUM! | #NUM! | -1.2436 | -0.6462 | #NUM! | -2.7691 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.054 | 0.058 | 0.052 | 0.054 |  |  | 0.028 | 0.028 | -0.104 | -0.023 |  |  | -0.008 | -0.004 | -0.01 | -0.008 | --- | --- | -0.03025 | -0.03025 | -0.16225 | -0.08125 | --- | --- | 2.8238 | 5.6477 | 12.1168 | 7.5847 | --- | --- | 7.0432 | 3.9066 | 6.1874 | 1.8407 | 1.0381 | 1.7312 | 2.4946 | 2.0261 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.054 | 0.054 | 0.058 | 0.057 |  |  | -0.011 | 0.08 | 0.046 | -0.069 |  |  | -0.008 | -0.008 | -0.004 | -0.005 | --- | --- | -0.06925 | 0.02175 | -0.01225 | -0.12725 | --- | --- | 6.4645 | -2.0304 | 2.2871 | 19.006 | --- | --- | 6.4318 | 9.0719 | #NUM! | #NUM! | 1.8663 | #NUM! | 0.82727 | 2.9448 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.112 | 0.112 | 0.12 | 0.13 |  |  | 0.302 | 0.441 | 0.012 | 0.306 |  |  | 0.05 | 0.05 | 0.058 | 0.068 | --- | --- | 0.24375 | 0.38275 | -0.04625 | 0.24775 | --- | --- | 3.6406 | 5.7167 | -0.59551 | 2.7209 | --- | --- | 2.8707 | 2.6287 | #NUM! | #NUM! | 1.2922 | 1.7434 | #NUM! | 1.001 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.128 | 0.142 | 0.136 | 0.153 |  |  | 0.681 | 0.592 | 0.4 | 0.396 |  |  | 0.066 | 0.08 | 0.074 | 0.091 | --- | --- | 0.62275 | 0.53375 | 0.34175 | 0.33775 | --- | --- | 7.0465 | 4.9825 | 3.4489 | 2.7718 | --- | --- | 4.5624 | 1.8968 | 4.2802 | 1.5085 | 1.9525 | 1.6059 | 1.2381 | 1.0195 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.11 | 0.12 | 0.121 | 0.125 |  |  | 0.057 | -0.029 | 0.075 | 0.121 |  |  | 0.048 | 0.058 | 0.059 | 0.063 | --- | --- | -0.00125 | -0.08725 | 0.01675 | 0.06275 | --- | --- | -0.019448 | -1.1234 | 0.21201 | 0.74383 | --- | --- | -0.046754 | 0.78569 | #NUM! | #NUM! | #NUM! | #NUM! | -1.5511 | -0.29594 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.127 | 0.136 | 0.14 | 0.143 |  |  | 0.018 | -0.035 | -0.085 | 0.058 |  |  | 0.065 | 0.074 | 0.078 | 0.081 | --- | --- | -0.04025 | -0.09325 | -0.14325 | -0.00025 | --- | --- | -0.46244 | -0.94107 | -1.3715 | -0.0023049 | --- | --- | -0.69433 | 0.59221 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.086 | 0.098 | 0.088 | 0.1 |  |  | 0.224 | 0.023 | 0.122 | 0.098 |  |  | 0.024 | 0.036 | 0.026 | 0.038 | --- | --- | 0.16575 | -0.03525 | 0.06375 | 0.03975 | --- | --- | 5.1576 | -0.73124 | 1.8311 | 0.78119 | --- | --- | 1.7597 | 2.4975 | #NUM! | #NUM! | 1.6405 | #NUM! | 0.60491 | -0.24694 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.089 | 0.09 | 0.093 | 0.097 |  |  | 0.149 | 0.433 | 0.202 | 0.472 |  |  | 0.027 | 0.028 | 0.031 | 0.035 | --- | --- | 0.09075 | 0.37475 | 0.14375 | 0.41375 | --- | --- | 2.5101 | 9.9951 | 3.463 | #NUM! | --- | --- | #NUM! | #NUM! | #NUM! | 2.0606 | 0.92031 | 2.3021 | 1.2421 | --- | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.113 | 0.108 | 0.109 | 0.124 |  |  | 0.15 | 0.091 | 0.006 | 0.206 |  |  | 0.051 | 0.046 | 0.047 | 0.062 | --- | --- | 0.09175 | 0.03275 | -0.05225 | 0.14775 | --- | --- | 1.3435 | 0.53169 | -0.83022 | #NUM! | --- | --- | #NUM! | #NUM! | #NUM! | #NUM! | 0.29528 | -0.6317 | #NUM! | --- | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.123 | 0.121 | 0.123 | 0.136 |  |  | 0.171 | 0.086 | -0.017 | 0.084 |  |  | 0.061 | 0.059 | 0.061 | 0.074 | --- | --- | 0.11275 | 0.02775 | -0.07525 | 0.02575 | --- | --- | 1.3804 | 0.35125 | -0.92125 | #NUM! | --- | --- | #NUM! | #NUM! | #NUM! | #NUM! | 0.32234 | -1.0463 | #NUM! | --- | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.128 | 0.117 | 0.125 | 0.13 |  |  | -0.108 | -0.051 | 0.096 | -0.058 |  |  | 0.066 | 0.055 | 0.063 | 0.068 | --- | --- | -0.16625 | -0.10925 | 0.03775 | -0.11625 | --- | --- | -1.8811 | -1.4834 | 0.44749 | #NUM! | --- | --- | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | -0.80411 | --- | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.13 | 0.124 | 0.132 | 0.124 |  |  | 0.006 | -0.041 | 0.122 | -0.227 |  |  | 0.068 | 0.062 | 0.07 | 0.062 | --- | --- | -0.05225 | -0.09925 | 0.06375 | -0.28525 | --- | --- | -0.57383 | -1.1955 | 0.68012 | #NUM! | --- | --- | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | -0.38549 | --- | --- | ---