## OD600 reference point
| LUDOX-HS40 | H2O
Replicate 1 | 0.0428 | 0.0318 | Enter Abs600 absorbance measurements into blue cells
Replicate 2 | 0.0422 | 0.0315 | Gold cells are calculated
Replicate 3 | 0.0436 | 0.03
Replicate 4 | 0.0442 | 0.0322
Arith. Mean | 0.0432 | 0.031375
Corrected Abs600 | 0.011825 |  | Corrected value is particle-only contribution
Reference OD600 | 0.0425 |  | Reference value is for 100uL of LUDOX-HS40 in a well of a standard 96-well flat-bottom plate
OD600/Abs600 | 3.5941 |  | Corrected value = scaling factor * measured value

## Fluorescein standard curve
uM Fluorescein | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625 | 0.78125 | 0.39062 | 0.19531 | 0.097656 | 0.048828 | 0
Replicate 1 | 15772 | 10810 | 6341 | 3403 | 1771 | 911 | 455 | 228 | 115 | 59 | 29 | 1 | Enter fluorescence measurements into blue cells
Replicate 2 | 16240 | 11472 | 5327 | 3039 | 1804 | 923 | 475 | 242 | 118 | 60 | 30 | 1 | Gold cells are calculated
Replicate 3 | 15466 | 11317 | 6707 | 3691 | 1838 | 939 | 468 | 235 | 121 | 59 | 29 | 1
Replicate 4 | 16597 | 11176 | 6549 | 3468 | 1851 | 931 | 464 | 233 | 116 | 57 | 29 | 0 | Values measured are fluorescence from 100uL of X uM fluorescein solution
Arith. Mean | 16018.75 | 11193.75 | 6231 | 3400.25 | 1816 | 926 | 465.5 | 234.5 | 117.5 | 58.75 | 29.25 | 0.75
Arith. Std.Dev. | 499.9142 | 282.9563 | 621.025 | 270.5764 | 35.9537 | 11.9443 | 8.3467 | 5.8023 | 2.6458 | 1.2583 | 0.5 | 0.5
Values should form a straight line on both linear and log scale
Slope should be 1:1
Common problems:
* Consistent pipetting error --> log graph is a straight line but not 1:1 slope
* Oversaturated detector --> low concentrations linear, but high concentrations saturate or fall
uM Fluorescein/a.u. | 50 | 25 | 12.5 | 6.25 | 3.125 | 1.5625 | 0.78125 | 0.39062 | 0.19531 | 0.097656 | 0.048828
Mean um Fluorescein/a.u. | 0.0031213 | 0.0022334 | 0.0020061 | 0.0018381 | 0.0017208 | 0.0016874 | 0.0016783 | 0.0016658 | 0.0016622 | 0.0016622 | 0.0016693
Mean of med-high levels: |  | 0.0018972
Final scaling level determined from medium-high points likely to be less impacted by saturation or pipetting error
If needed, you can shift which points are used, but it is likely better to correct instrument settings and protocol.

## Raw Plate Reader Measurements
Raw Plate Readings |  | If you followed the recommended plate layout:
Copy fluorescence and Abs600 measurements from your plate reader into blue cells
They will automatically propagate into the correct locations in the Fluorescence Measurement Sheet
Fluorescence Raw Readings: | [9 empty] | Abs600 Raw Readings:
Hour 0: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 0: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 16 | 27 | 92 | 47 | 16 | 25 | 18 | 15 | 14 | Colony 1, Replicate 1 | 0.0671 | 0.0679 | 0.0725 | 0.0694 | 0.0704 | 0.0633 | 0.0626 | 0.0653 | 0.0401
Colony 1, Replicate 2 | 16 | 29 | 84 | 47 | 16 | 25 | 18 | 17 | 15 | Colony 1, Replicate 2 | 0.0652 | 0.0686 | 0.0699 | 0.0696 | 0.0712 | 0.0644 | 0.0626 | 0.0664 | 0.0424
Colony 1, Replicate 3 | 16 | 30 | 99 | 50 | 17 | 27 | 19 | 17 | 17 | Colony 1, Replicate 3 | 0.0646 | 0.0691 | 0.0721 | 0.0696 | 0.0693 | 0.0634 | 0.0635 | 0.0677 | 0.0427
Colony 1, Replicate 4 | 16 | 29 | 99 | 49 | 18 | 27 | 19 | 17 | 17 | Colony 1, Replicate 4 | 0.067 | 0.0708 | 0.0746 | 0.0709 | 0.0732 | 0.0652 | 0.0657 | 0.0697 | 0.0425
Colony 2, Replicate 1 | 16 | 28 | 82 | 41 | 16 | 27 | 18 | 16 | 16 | Colony 2, Replicate 1 | 0.0784 | 0.0716 | 0.0794 | 0.0788 | 0.079 | 0.075 | 0.0749 | 0.0945 | 0.0378
Colony 2, Replicate 2 | 16 | 28 | 79 | 42 | 16 | 26 | 18 | 16 | 16 | Colony 2, Replicate 2 | 0.0745 | 0.0708 | 0.0768 | 0.0752 | 0.0705 | 0.071 | 0.072 | 0.0791 | 0.0361
Colony 2, Replicate 3 | 15 | 26 | 81 | 40 | 16 | 26 | 18 | 16 | 15 | Colony 2, Replicate 3 | 0.076 | 0.0726 | 0.079 | 0.0743 | 0.0731 | 0.0728 | 0.0726 | 0.0794 | 0.0366
Colony 2, Replicate 4 | 15 | 26 | 79 | 41 | 15 | 25 | 18 | 15 | 15 | Colony 2, Replicate 4 | 0.074 | 0.0665 | 0.0707 | 0.0697 | 0.0701 | 0.075 | 0.0685 | 0.0739 | 0.0398
Hour 2: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 2: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 14 | 41 | 73 | 75 | 14 | 48 | 28 | 15 | 15 | Colony 1, Replicate 1 | 0.1139 | 0.0895 | 0.0623 | 0.1052 | 0.1182 | 0.0833 | 0.1082 | 0.1398 | 0.0395
Colony 1, Replicate 2 | 13 | 43 | 73 | 74 | 16 | 50 | 29 | 15 | 14 | Colony 1, Replicate 2 | 0.1158 | 0.0884 | 0.062 | 0.1094 | 0.1195 | 0.0874 | 0.1069 | 0.1376 | 0.0386
Colony 1, Replicate 3 | 13 | 47 | 79 | 77 | 17 | 50 | 31 | 15 | 14 | Colony 1, Replicate 3 | 0.1157 | 0.0879 | 0.0618 | 0.1064 | 0.1131 | 0.0825 | 0.1035 | 0.1346 | 0.0392
Colony 1, Replicate 4 | 15 | 45 | 81 | 77 | 17 | 52 | 30 | 16 | 15 | Colony 1, Replicate 4 | 0.107 | 0.07 | 0.063 | 0.1125 | 0.1223 | 0.0881 | 0.1088 | 0.1423 | 0.04
Colony 2, Replicate 1 | 14 | 43 | 66 | 68 | 15 | 45 | 27 | 15 | 15 | Colony 2, Replicate 1 | 0.1216 | 0.1003 | 0.0725 | 0.1244 | 0.1208 | 0.0971 | 0.1205 | 0.1278 | 0.0386
Colony 2, Replicate 2 | 15 | 43 | 66 | 69 | 16 | 45 | 28 | 15 | 14 | Colony 2, Replicate 2 | 0.1184 | 0.1006 | 0.0668 | 0.1168 | 0.1167 | 0.0917 | 0.1124 | 0.1159 | 0.0362
Colony 2, Replicate 3 | 13 | 43 | 67 | 68 | 16 | 45 | 28 | 14 | 14 | Colony 2, Replicate 3 | 0.121 | 0.0971 | 0.0654 | 0.1201 | 0.1137 | 0.0925 | 0.1159 | 0.1201 | 0.0355
Colony 2, Replicate 4 | 15 | 42 | 62 | 66 | 16 | 43 | 27 | 15 | 15 | Colony 2, Replicate 4 | 0.1124 | 0.0945 | 0.0632 | 0.1122 | 0.1085 | 0.0858 | 0.1069 | 0.1123 | 0.0378
Hour 4: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 4: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 20 | 128 | 133 | 189 | 22 | 141 | 51 | 20 | 17 | Colony 1, Replicate 1 | 0.2298 | 0.1933 | 0.0682 | 0.2228 | 0.257 | 0.132 | 0.2264 | 0.258 | 0.0384
Colony 1, Replicate 2 | 17 | 131 | 119 | 177 | 21 | 143 | 55 | 20 | 16 | Colony 1, Replicate 2 | 0.2183 | 0.1886 | 0.0652 | 0.2229 | 0.2402 | 0.1352 | 0.2313 | 0.2491 | 0.0412
Colony 1, Replicate 3 | 19 | 134 | 139 | 200 | 22 | 144 | 55 | 21 | 17 | Colony 1, Replicate 3 | 0.219 | 0.1758 | 0.0672 | 0.2004 | 0.2335 | 0.1228 | 0.2184 | 0.2469 | 0.0427
Colony 1, Replicate 4 | 20 | 133 | 138 | 196 | 22 | 150 | 55 | 20 | 17 | Colony 1, Replicate 4 | 0.2117 | 0.2012 | 0.0682 | 0.2306 | 0.2544 | 0.1351 | 0.2199 | 0.2515 | 0.0425
Colony 2, Replicate 1 | 19 | 105 | 99 | 163 | 22 | 139 | 44 | 20 | 18 | Colony 2, Replicate 1 | 0.2807 | 0.2041 | 0.0776 | 0.2445 | 0.2684 | 0.1746 | 0.2232 | 0.2904 | 0.0379
Colony 2, Replicate 2 | 20 | 99 | 96 | 160 | 22 | 139 | 45 | 22 | 18 | Colony 2, Replicate 2 | 0.2619 | 0.1959 | 0.0701 | 0.2321 | 0.2515 | 0.1662 | 0.2006 | 0.26 | 0.0369
Colony 2, Replicate 3 | 18 | 103 | 96 | 159 | 23 | 137 | 43 | 21 | 17 | Colony 2, Replicate 3 | 0.2728 | 0.2008 | 0.0703 | 0.2316 | 0.2542 | 0.159 | 0.2104 | 0.2541 | 0.0373
Colony 2, Replicate 4 | 20 | 101 | 93 | 155 | 22 | 131 | 42 | 20 | 17 | Colony 2, Replicate 4 | 0.2559 | 0.1905 | 0.0635 | 0.2218 | 0.242 | 0.1498 | 0.1994 | 0.2482 | 0.0402
Hour 6: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank) | Hour 6: | Neg. Control | Pos. Control | Device 1 | Device 2 | Device 3 | Device 4 | Device 5 | Device 6 | LB + Chlor (blank)
Colony 1, Replicate 1 | 17 | 124 | 125 | 177 | 20 | 128 | 42 | 17 | 13 | Colony 1, Replicate 1 | 0.3454 | 0.2813 | 0.0712 | 0.2683 | 0.3073 | 0.2284 | 0.301 | 0.3262 | 0.0403
Colony 1, Replicate 2 | 16 | 126 | 121 | 175 | 20 | 133 | 44 | 19 | 13 | Colony 1, Replicate 2 | 0.3352 | 0.2822 | 0.0712 | 0.2648 | 0.3037 | 0.2231 | 0.3001 | 0.3275 | 0.0417
Colony 1, Replicate 3 | 16 | 136 | 132 | 182 | 20 | 138 | 46 | 19 | 15 | Colony 1, Replicate 3 | 0.34 | 0.2814 | 0.0723 | 0.2631 | 0.2895 | 0.2077 | 0.2829 | 0.3065 | 0.0417
Colony 1, Replicate 4 | 18 | 138 | 130 | 190 | 20 | 141 | 46 | 18 | 14 | Colony 1, Replicate 4 | 0.3211 | 0.2899 | 0.0729 | 0.2861 | 0.3092 | 0.2317 | 0.3033 | 0.3301 | 0.0454
Colony 2, Replicate 1 | 17 | 123 | 88 | 163 | 19 | 142 | 37 | 17 | 14 | Colony 2, Replicate 1 | 0.3626 | 0.3011 | 0.0714 | 0.3512 | 0.3716 | 0.3092 | 0.2743 | 0.4117 | 0.0381
Colony 2, Replicate 2 | 15 | 115 | 87 | 170 | 20 | 138 | 35 | 17 | 13 | Colony 2, Replicate 2 | 0.3601 | 0.3087 | 0.0674 | 0.345 | 0.3474 | 0.3063 | 0.2718 | 0.3814 | 0.0359
Colony 2, Replicate 3 | 16 | 121 | 90 | 174 | 19 | 142 | 35 | 17 | 12 | Colony 2, Replicate 3 | 0.3534 | 0.2982 | 0.0684 | 0.3431 | 0.3446 | 0.2912 | 0.2699 | 0.3921 | 0.0358
Colony 2, Replicate 4 | 17 | 116 | 83 | 160 | 19 | 139 | 34 | 18 | 14 | Colony 2, Replicate 4 | 0.3384 | 0.2831 | 0.0647 | 0.3215 | 0.3314 | 0.2736 | 0.2667 | 0.3602 | 0.038
Plate pattern:
| A1 | A2 | A3 | A4 | A5 | A6 | A7 | A8 | A9
| B1 | B2 | B3 | B4 | B5 | B6 | B7 | B8 | B9
| C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9
| D1 | D2 | D3 | D4 | D5 | D6 | D7 | D8 | D9
| E1 | E2 | E3 | E4 | E5 | E6 | E7 | E8 | E9
| F1 | F2 | F3 | F4 | F5 | F6 | F7 | F8 | F9
| G1 | G2 | G3 | G4 | G5 | G6 | G7 | G8 | G9
| H1 | H2 | H3 | H4 | H5 | H6 | H7 | H8 | H9

## Fluorescence Measurement
Unit Scaling Factors: | These are imported from the prior two sheets | [5 empty] | Enter fluorescence and Abs600 measurements into blue cells on "Raw Plate Reader Measurements" | [34 empty] | 0 | 2 | 4 | 6 |  | 0 | 2 | 4 | 6 | 0 | 2 | 4 | 6 |  | 0h | 2h | 4h | 6h |  | 0h | 2h | 4h | 6h
OD600/Abs600 | 3.5941 | [5 empty] | They will be copied into the green cells on this sheet. | [27 empty] | Negative Control (Colony 1) | [6 empty] | 0.025175 | 0.071975 | 0.18788 | 0.30347 | Negative Control (Colony 1) | 0.25 | -1.75 | 4.25 | 1.25 | 0.0052419 | -0.012834 | 0.011941 | 0.0021742 | (-) Control | 0.0052419 | -0.012834 | 0.011941 | 0.0021742
uM Fluorescein/a.u. | 0.0018972 | [5 empty] | Gold cells are calculated | [27 empty] | Negative Control (Colony 2) | [6 empty] | 0.036475 | 0.079675 | 0.23878 | 0.32067 |  | 0.25 | -1.75 | 3.25 | 1.25 | 0.0036179 | -0.011594 | 0.0071847 | 0.0020576 | [5 empty] | (-) Control | 0.0036179 | -0.011594 | 0.0071847 | 0.0020576
[7 empty] | If you have more replicates, unhide the extra columns | [27 empty] | Positive Control (Colony 1) | [6 empty] | 0.025975 | 0.047575 | 0.15138 | 0.23938 | Positive Control (Colony 1) | 11.25 | 25.25 | 112.25 | 108.25 | 0.22862 | 0.28015 | 0.39142 | 0.23871 | (+) Control | 0.22862 | 0.28015 | 0.39142 | 0.23871
[35 empty] | Positive Control (Colony 2) | [6 empty] | 0.029675 | 0.058375 | 0.16218 | 0.25917 |  | 12.25 | 27.25 | 89.25 | 107.25 | 0.2179 | 0.24641 | 0.2905 | 0.21843 | [5 empty] | (+) Control | 0.2179 | 0.24641 | 0.2905 | 0.21843
Experimental Values: | Raw Abs600 | [5 empty] | Raw Fluorescence | [27 empty] | Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | [6 empty] | 0.030575 | 0.020375 | 0.026275 | 0.029275 | Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 76.25 | 57.25 | 117.25 | 109.25 | 1.3164 | 1.4832 | 2.3555 | 1.9699 | Device 1 | 1.3164 | 1.4832 | 2.3555 | 1.9699
Sample set: | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | [22 empty] | Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | [6 empty] | 0.037475 | 0.030575 | 0.035675 | 0.029475 |  | 66.25 | 50.25 | 83.25 | 72.25 | 0.93317 | 0.86753 | 1.2318 | 1.2939 | [5 empty] | Device 1 | 0.93317 | 0.86753 | 1.2318 | 1.2939
Blank media | 0.0401 | 0.0424 | 0.0427 | 0.0425 |  |  | 14 | 15 | 17 | 17 | [24 empty] | Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | [6 empty] | 0.027475 | 0.063275 | 0.18088 | 0.22637 | Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 31.25 | 59.25 | 173.25 | 161.25 | 0.60038 | 0.49428 | 0.5056 | 0.376 | Device 2 | 0.60038 | 0.49428 | 0.5056 | 0.376
Blank mean: | 0.041925 | [5 empty] | 15.75 | [5 empty] | OD - Background | [5 empty] | Fluorescence - Background | [5 empty] | uM Fluorescein / OD600 | [5 empty] | Summary Statistics |  |  |  | Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | Ln uM FITC / OD600 | [5 empty] | 0.036875 | 0.082475 | 0.20258 | 0.30928 |  | 25.25 | 52.25 | 147.25 | 147.25 | 0.36145 | 0.33441 | 0.38369 | 0.25132 | [5 empty] | Device 2 | 0.36145 | 0.33441 | 0.38369 | 0.25132
Hour 0: | [12 empty] | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | Arith. Mean | Arith. Std.Dev. | Geo. Mean | Geo. Std. Dev. | Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | Replicate 1 | Replicate 2 | Replicate 3 | Replicate 4 | Replicate 5 | Replicate 6 | 0.028475 | 0.076275 | 0.21508 | 0.26538 | Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.25 | -1.75 | 6.25 | 4.25 | 0.0046344 | -0.012111 | 0.015339 | 0.0084536 | Device 3 | 0.0046344 | -0.012111 | 0.015339 | 0.0084536
Negative Control (Colony 1) | 0.0671 | 0.0652 | 0.0646 | 0.067 |  |  | 16 | 16 | 16 | 16 |  |  | 0.025175 | 0.023275 | 0.022675 | 0.025075 | --- | --- | 0.25 | 0.25 | 0.25 | 0.25 | --- | --- | 0.0052419 | 0.0056698 | 0.0058198 | 0.0052628 | --- | --- | 0.0054985 | 0.00029097 | 0.0054928 | 1.0542 | Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | -5.2511 | -5.1726 | -5.1465 | -5.2471 | --- | --- | 0.037075 | 0.078875 | 0.22648 | 0.32967 |  | 0.25 | -0.75 | 6.25 | 3.25 | 0.0035594 | -0.0050192 | 0.014567 | 0.0052037 | [5 empty] | Device 3 | 0.0035594 | -0.0050192 | 0.014567 | 0.0052037
Negative Control (Colony 2) | 0.0784 | 0.0745 | 0.076 | 0.074 |  |  | 16 | 16 | 15 | 15 |  |  | 0.036475 | 0.032575 | 0.034075 | 0.032075 | --- | --- | 0.25 | 0.25 | -0.75 | -0.75 | --- | --- | 0.0036179 | 0.0040511 | -0.011618 | -0.012343 | --- | --- | -0.004073 | 0.0091373 | #NUM! | #NUM! | Test Device 4: J23101+I13504 (Colony 1) | -5.6219 | -5.5088 | #NUM! | #NUM! | --- | --- | 0.021375 | 0.041375 | 0.090075 | 0.18648 | Test Device 4: J23101+I13504 (Colony 1) | 9.25 | 32.25 | 125.25 | 112.25 | 0.22843 | 0.41144 | 0.73399 | 0.31775 | Device 4 | 0.22843 | 0.41144 | 0.73399 | 0.31775
Positive Control (Colony 1) | 0.0679 | 0.0686 | 0.0691 | 0.0708 |  |  | 27 | 29 | 30 | 29 |  |  | 0.025975 | 0.026675 | 0.027175 | 0.028875 | --- | --- | 11.25 | 13.25 | 14.25 | 13.25 | --- | --- | 0.22862 | 0.2622 | 0.2768 | 0.24222 | --- | --- | 0.25246 | 0.021294 | 0.25178 | 1.0882 | Test Device 4: J23101+I13504 (Colony 2) | -1.4757 | -1.3387 | -1.2845 | -1.4179 | --- | --- | 0.033075 | 0.055175 | 0.13268 | 0.26727 |  | 11.25 | 29.25 | 123.25 | 126.25 | 0.17954 | 0.27983 | 0.49036 | 0.24934 | [5 empty] | Device 4 | 0.17954 | 0.27983 | 0.49036 | 0.24934
Positive Control (Colony 2) | 0.0716 | 0.0708 | 0.0726 | 0.0665 |  |  | 28 | 28 | 26 | 26 |  |  | 0.029675 | 0.028875 | 0.030675 | 0.024575 | --- | --- | 12.25 | 12.25 | 10.25 | 10.25 | --- | --- | 0.2179 | 0.22394 | 0.17638 | 0.22016 | --- | --- | 0.2096 | 0.022283 | 0.20864 | 1.1191 | Test Device 5: J23106+I13504 (Colony 1) | -1.5237 | -1.4964 | -1.7351 | -1.5134 | --- | --- | 0.020675 | 0.066275 | 0.18448 | 0.25907 | Test Device 5: J23106+I13504 (Colony 1) | 2.25 | 12.25 | 35.25 | 26.25 | 0.057445 | 0.097567 | 0.10086 | 0.053483 | Device 5 | 0.057445 | 0.097567 | 0.10086 | 0.053483
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.0725 | 0.0699 | 0.0721 | 0.0746 |  |  | 92 | 84 | 99 | 99 |  |  | 0.030575 | 0.027975 | 0.030175 | 0.032675 | --- | --- | 76.25 | 68.25 | 83.25 | 83.25 | --- | --- | 1.3164 | 1.2878 | 1.4563 | 1.3449 | --- | --- | 1.3513 | 0.073751 | 1.3499 | 1.0551 | Test Device 5: J23106+I13504 (Colony 2) | 0.2749 | 0.25293 | 0.3759 | 0.2963 | --- | --- | 0.032975 | 0.078575 | 0.18128 | 0.23237 |  | 2.25 | 11.25 | 28.25 | 21.25 | 0.036017 | 0.075576 | 0.082261 | 0.048271 | [5 empty] | Device 5 | 0.036017 | 0.075576 | 0.082261 | 0.048271
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.0794 | 0.0768 | 0.079 | 0.0707 |  |  | 82 | 79 | 81 | 79 |  |  | 0.037475 | 0.034875 | 0.037075 | 0.028775 | --- | --- | 66.25 | 63.25 | 65.25 | 63.25 | --- | --- | 0.93317 | 0.95733 | 0.929 | 1.1603 | --- | --- | 0.99494 | 0.11093 | 0.9906 | 1.1121 | Test Device 6: J23117+I13504 (Colony 1) | -0.069172 | -0.043609 | -0.07365 | 0.14865 | --- | --- | 0.023375 | 0.097875 | 0.21608 | 0.28427 | Test Device 6: J23117+I13504 (Colony 1) | -0.75 | -0.75 | 4.25 | 1.25 | -0.016937 | -0.0040449 | 0.010382 | 0.0023211 | Device 6 | -0.016937 | -0.0040449 | 0.010382 | 0.0023211
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.0694 | 0.0696 | 0.0696 | 0.0709 |  |  | 47 | 47 | 50 | 49 |  |  | 0.027475 | 0.027675 | 0.027675 | 0.028975 | --- | --- | 31.25 | 31.25 | 34.25 | 33.25 | --- | --- | 0.60038 | 0.59604 | 0.65326 | 0.60574 | --- | --- | 0.61386 | 0.026569 | 0.61344 | 1.0434 | Test Device 6: J23117+I13504 (Colony 2) | -0.51019 | -0.51744 | -0.42578 | -0.50131 | --- | --- | 0.052575 | 0.085875 | 0.24848 | 0.36978 |  | 0.25 | -0.75 | 4.25 | 1.25 | 0.00251 | -0.0046101 | 0.0090286 | 0.0017844 | [5 empty] | Device 6 | 0.00251 | -0.0046101 | 0.0090286 | 0.0017844
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.0788 | 0.0752 | 0.0743 | 0.0697 |  |  | 41 | 42 | 40 | 41 |  |  | 0.036875 | 0.033275 | 0.032375 | 0.027775 | --- | --- | 25.25 | 26.25 | 24.25 | 25.25 | --- | --- | 0.36145 | 0.41641 | 0.39538 | 0.47987 | --- | --- | 0.41328 | 0.049836 | 0.41108 | 1.1259 |  | -1.0176 | -0.87607 | -0.9279 | -0.73424 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.0704 | 0.0712 | 0.0693 | 0.0732 |  |  | 16 | 16 | 17 | 18 |  |  | 0.028475 | 0.029275 | 0.027375 | 0.031275 | --- | --- | 0.25 | 0.25 | 1.25 | 2.25 | --- | --- | 0.0046344 | 0.0045077 | 0.024103 | 0.037975 | --- | --- | 0.017805 | 0.016297 | 0.011759 | 3.025 |  | -5.3743 | -5.402 | -3.7254 | -3.2708 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.079 | 0.0705 | 0.0731 | 0.0701 |  |  | 16 | 16 | 16 | 15 |  |  | 0.037075 | 0.028575 | 0.031175 | 0.028175 | --- | --- | 0.25 | 0.25 | 0.25 | -0.75 | --- | --- | 0.0035594 | 0.0046182 | 0.004233 | -0.014051 | --- | --- | -0.00041016 | 0.0091045 | #NUM! | #NUM! |  | -5.6382 | -5.3778 | -5.4648 | #NUM! | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.0633 | 0.0644 | 0.0634 | 0.0652 |  |  | 25 | 25 | 27 | 27 |  |  | 0.021375 | 0.022475 | 0.021475 | 0.023275 | --- | --- | 9.25 | 9.25 | 11.25 | 11.25 | --- | --- | 0.22843 | 0.21725 | 0.27652 | 0.25514 | --- | --- | 0.24434 | 0.026706 | 0.24325 | 1.1148 |  | -1.4765 | -1.5267 | -1.2855 | -1.3659 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.075 | 0.071 | 0.0728 | 0.075 |  |  | 27 | 26 | 26 | 25 |  |  | 0.033075 | 0.029075 | 0.030875 | 0.033075 | --- | --- | 11.25 | 10.25 | 10.25 | 9.25 | --- | --- | 0.17954 | 0.18609 | 0.17524 | 0.14762 | --- | --- | 0.17212 | 0.016931 | 0.17146 | 1.1083 |  | -1.7173 | -1.6815 | -1.7416 | -1.9131 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.0626 | 0.0626 | 0.0635 | 0.0657 |  |  | 18 | 18 | 19 | 19 |  |  | 0.020675 | 0.020675 | 0.021575 | 0.023775 | --- | --- | 2.25 | 2.25 | 3.25 | 3.25 | --- | --- | 0.057445 | 0.057445 | 0.079515 | 0.072157 | --- | --- | 0.06664 | 0.011035 | 0.065964 | 1.1788 |  | -2.8569 | -2.8569 | -2.5318 | -2.6289 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.0749 | 0.072 | 0.0726 | 0.0685 |  |  | 18 | 18 | 18 | 18 |  |  | 0.032975 | 0.030075 | 0.030675 | 0.026575 | --- | --- | 2.25 | 2.25 | 2.25 | 2.25 | --- | --- | 0.036017 | 0.03949 | 0.038718 | 0.044691 | --- | --- | 0.039729 | 0.0036277 | 0.039608 | 1.0938 |  | -3.3238 | -3.2317 | -3.2515 | -3.108 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.0653 | 0.0664 | 0.0677 | 0.0697 |  |  | 15 | 17 | 17 | 17 |  |  | 0.023375 | 0.024475 | 0.025775 | 0.027775 | --- | --- | -0.75 | 1.25 | 1.25 | 1.25 | --- | --- | -0.016937 | 0.026959 | 0.025599 | 0.023756 | --- | --- | 0.014844 | 0.021228 | #NUM! | #NUM! |  | #NUM! | -3.6134 | -3.6652 | -3.7399 | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.0945 | 0.0791 | 0.0794 | 0.0739 |  |  | 16 | 16 | 16 | 15 |  |  | 0.052575 | 0.037175 | 0.037475 | 0.031975 | --- | --- | 0.25 | 0.25 | 0.25 | -0.75 | --- | --- | 0.00251 | 0.0035498 | 0.0035214 | -0.012381 | --- | --- | -0.00070002 | 0.0078025 | #NUM! | #NUM! |  | -5.9875 | -5.6409 | -5.6489 | #NUM! | --- | ---
Hour 2:
Negative Control (Colony 1) | 0.1139 | 0.1158 | 0.1157 | 0.107 |  |  | 14 | 13 | 13 | 15 |  |  | 0.071975 | 0.073875 | 0.073775 | 0.065075 | --- | --- | -1.75 | -2.75 | -2.75 | -0.75 | --- | --- | -0.012834 | -0.019649 | -0.019676 | -0.0060836 | --- | --- | -0.014561 | 0.0065039 | #NUM! | #NUM! |  | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Negative Control (Colony 2) | 0.1216 | 0.1184 | 0.121 | 0.1124 |  |  | 14 | 15 | 13 | 15 |  |  | 0.079675 | 0.076475 | 0.079075 | 0.070475 | --- | --- | -1.75 | -0.75 | -2.75 | -0.75 | --- | --- | -0.011594 | -0.0051767 | -0.018357 | -0.0056175 | --- | --- | -0.010186 | 0.0061837 | #NUM! | #NUM! |  | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Positive Control (Colony 1) | 0.0895 | 0.0884 | 0.0879 | 0.07 |  |  | 41 | 43 | 47 | 45 |  |  | 0.047575 | 0.046475 | 0.045975 | 0.028075 | --- | --- | 25.25 | 27.25 | 31.25 | 29.25 | --- | --- | 0.28015 | 0.3095 | 0.35879 | 0.54995 | --- | --- | 0.3746 | 0.12132 | 0.36166 | 1.3462 |  | -1.2724 | -1.1728 | -1.025 | -0.59793 | --- | ---
Positive Control (Colony 2) | 0.1003 | 0.1006 | 0.0971 | 0.0945 |  |  | 43 | 43 | 43 | 42 |  |  | 0.058375 | 0.058675 | 0.055175 | 0.052575 | --- | --- | 27.25 | 27.25 | 27.25 | 26.25 | --- | --- | 0.24641 | 0.24515 | 0.2607 | 0.26355 | --- | --- | 0.25395 | 0.0095235 | 0.25382 | 1.0382 |  | -1.4008 | -1.4059 | -1.3444 | -1.3335 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.0623 | 0.062 | 0.0618 | 0.063 |  |  | 73 | 73 | 79 | 81 |  |  | 0.020375 | 0.020075 | 0.019875 | 0.021075 | --- | --- | 57.25 | 57.25 | 63.25 | 65.25 | --- | --- | 1.4832 | 1.5053 | 1.6798 | 1.6343 | --- | --- | 1.5757 | 0.096244 | 1.5735 | 1.0629 |  | 0.39419 | 0.40902 | 0.5187 | 0.49121 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.0725 | 0.0668 | 0.0654 | 0.0632 |  |  | 66 | 66 | 67 | 62 |  |  | 0.030575 | 0.024875 | 0.023475 | 0.021275 | --- | --- | 50.25 | 50.25 | 51.25 | 46.25 | --- | --- | 0.86753 | 1.0663 | 1.1524 | 1.1475 | --- | --- | 1.0584 | 0.13326 | 1.0517 | 1.1425 |  | -0.14211 | 0.064214 | 0.14185 | 0.1376 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.1052 | 0.1094 | 0.1064 | 0.1125 |  |  | 75 | 74 | 77 | 77 |  |  | 0.063275 | 0.067475 | 0.064475 | 0.070575 | --- | --- | 59.25 | 58.25 | 61.25 | 61.25 | --- | --- | 0.49428 | 0.45569 | 0.50145 | 0.45811 | --- | --- | 0.47738 | 0.023853 | 0.47694 | 1.0512 |  | -0.70466 | -0.78595 | -0.69025 | -0.78065 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.1244 | 0.1168 | 0.1201 | 0.1122 |  |  | 68 | 69 | 68 | 66 |  |  | 0.082475 | 0.074875 | 0.078175 | 0.070275 | --- | --- | 52.25 | 53.25 | 52.25 | 50.25 | --- | --- | 0.33441 | 0.3754 | 0.3528 | 0.37744 | --- | --- | 0.36001 | 0.020397 | 0.35958 | 1.0589 |  | -1.0954 | -0.97976 | -1.0418 | -0.97434 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.1182 | 0.1195 | 0.1131 | 0.1223 |  |  | 14 | 16 | 17 | 17 |  |  | 0.076275 | 0.077575 | 0.071175 | 0.080375 | --- | --- | -1.75 | 0.25 | 1.25 | 1.25 | --- | --- | -0.012111 | 0.0017011 | 0.0092704 | 0.0082093 | --- | --- | 0.0017675 | 0.0098387 | #NUM! | #NUM! |  | #NUM! | -6.3765 | -4.6809 | -4.8025 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.1208 | 0.1167 | 0.1137 | 0.1085 |  |  | 15 | 16 | 16 | 16 |  |  | 0.078875 | 0.074775 | 0.071775 | 0.066575 | --- | --- | -0.75 | 0.25 | 0.25 | 0.25 | --- | --- | -0.0050192 | 0.0017648 | 0.0018386 | 0.0019822 | --- | --- | 0.00014159 | 0.0034417 | #NUM! | #NUM! |  | #NUM! | -6.3397 | -6.2988 | -6.2236 | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.0833 | 0.0874 | 0.0825 | 0.0881 |  |  | 48 | 50 | 50 | 52 |  |  | 0.041375 | 0.045475 | 0.040575 | 0.046175 | --- | --- | 32.25 | 34.25 | 34.25 | 36.25 | --- | --- | 0.41144 | 0.39756 | 0.44557 | 0.4144 | --- | --- | 0.41724 | 0.020262 | 0.41688 | 1.0491 |  | -0.88809 | -0.92241 | -0.8084 | -0.88093 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.0971 | 0.0917 | 0.0925 | 0.0858 |  |  | 45 | 45 | 45 | 43 |  |  | 0.055175 | 0.049775 | 0.050575 | 0.043875 | --- | --- | 29.25 | 29.25 | 29.25 | 27.25 | --- | --- | 0.27983 | 0.31019 | 0.30528 | 0.32784 | --- | --- | 0.30579 | 0.01983 | 0.3053 | 1.0678 |  | -1.2736 | -1.1706 | -1.1865 | -1.1152 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.1082 | 0.1069 | 0.1035 | 0.1088 |  |  | 28 | 29 | 31 | 30 |  |  | 0.066275 | 0.064975 | 0.061575 | 0.066875 | --- | --- | 12.25 | 13.25 | 15.25 | 14.25 | --- | --- | 0.097567 | 0.10764 | 0.13073 | 0.11248 | --- | --- | 0.1121 | 0.013885 | 0.11148 | 1.1294 |  | -2.3272 | -2.2289 | -2.0346 | -2.185 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.1205 | 0.1124 | 0.1159 | 0.1069 |  |  | 27 | 28 | 28 | 27 |  |  | 0.078575 | 0.070475 | 0.073975 | 0.064975 | --- | --- | 11.25 | 12.25 | 12.25 | 11.25 | --- | --- | 0.075576 | 0.091752 | 0.087411 | 0.091395 | --- | --- | 0.086533 | 0.0075654 | 0.086272 | 1.0952 |  | -2.5826 | -2.3887 | -2.4371 | -2.3926 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.1398 | 0.1376 | 0.1346 | 0.1423 |  |  | 15 | 15 | 15 | 16 |  |  | 0.097875 | 0.095675 | 0.092675 | 0.10038 | --- | --- | -0.75 | -0.75 | -0.75 | 0.25 | --- | --- | -0.0040449 | -0.0041379 | -0.0042718 | 0.0013147 | --- | --- | -0.002785 | 0.0027347 | #NUM! | #NUM! |  | #NUM! | #NUM! | #NUM! | -6.6341 | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.1278 | 0.1159 | 0.1201 | 0.1123 |  |  | 15 | 15 | 14 | 15 |  |  | 0.085875 | 0.073975 | 0.078175 | 0.070375 | --- | --- | -0.75 | -0.75 | -1.75 | -0.75 | --- | --- | -0.0046101 | -0.0053517 | -0.011816 | -0.0056255 | --- | --- | -0.0068509 | 0.003338 | #NUM! | #NUM! |  | #NUM! | #NUM! | #NUM! | #NUM! | --- | ---
Hour 4:
Negative Control (Colony 1) | 0.2298 | 0.2183 | 0.219 | 0.2117 |  |  | 20 | 17 | 19 | 20 |  |  | 0.18788 | 0.17638 | 0.17708 | 0.16978 | --- | --- | 4.25 | 1.25 | 3.25 | 4.25 | --- | --- | 0.011941 | 0.003741 | 0.0096881 | 0.013214 | --- | --- | 0.009646 | 0.0041979 | 0.0086961 | 1.7807 |  | -4.4278 | -5.5884 | -4.6369 | -4.3265 | --- | ---
Negative Control (Colony 2) | 0.2807 | 0.2619 | 0.2728 | 0.2559 |  |  | 19 | 20 | 18 | 20 |  |  | 0.23878 | 0.21998 | 0.23087 | 0.21398 | --- | --- | 3.25 | 4.25 | 2.25 | 4.25 | --- | --- | 0.0071847 | 0.010198 | 0.0051442 | 0.010484 | --- | --- | 0.0082529 | 0.002554 | 0.0079287 | 1.3991 |  | -4.9358 | -4.5855 | -5.2699 | -4.5579 | --- | ---
Positive Control (Colony 1) | 0.1933 | 0.1886 | 0.1758 | 0.2012 |  |  | 128 | 131 | 134 | 133 |  |  | 0.15138 | 0.14667 | 0.13388 | 0.15927 | --- | --- | 112.25 | 115.25 | 118.25 | 117.25 | --- | --- | 0.39142 | 0.41476 | 0.46625 | 0.38858 | --- | --- | 0.41525 | 0.035963 | 0.41413 | 1.0879 |  | -0.93797 | -0.88005 | -0.76304 | -0.94526 | --- | ---
Positive Control (Colony 2) | 0.2041 | 0.1959 | 0.2008 | 0.1905 |  |  | 105 | 99 | 103 | 101 |  |  | 0.16218 | 0.15398 | 0.15888 | 0.14858 | --- | --- | 89.25 | 83.25 | 87.25 | 85.25 | --- | --- | 0.2905 | 0.2854 | 0.28988 | 0.30287 | --- | --- | 0.29216 | 0.0074945 | 0.29209 | 1.0258 |  | -1.2362 | -1.2539 | -1.2383 | -1.1944 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.0682 | 0.0652 | 0.0672 | 0.0682 |  |  | 133 | 119 | 139 | 138 |  |  | 0.026275 | 0.023275 | 0.025275 | 0.026275 | --- | --- | 117.25 | 103.25 | 123.25 | 122.25 | --- | --- | 2.3555 | 2.3416 | 2.574 | 2.456 | --- | --- | 2.4318 | 0.10764 | 2.43 | 1.0448 |  | 0.85676 | 0.85084 | 0.94547 | 0.89852 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.0776 | 0.0701 | 0.0703 | 0.0635 |  |  | 99 | 96 | 96 | 93 |  |  | 0.035675 | 0.028175 | 0.028375 | 0.021575 | --- | --- | 83.25 | 80.25 | 80.25 | 77.25 | --- | --- | 1.2318 | 1.5035 | 1.4929 | 1.89 | --- | --- | 1.5295 | 0.27118 | 1.5119 | 1.1914 |  | 0.20846 | 0.40778 | 0.4007 | 0.63658 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.2228 | 0.2229 | 0.2004 | 0.2306 |  |  | 189 | 177 | 200 | 196 |  |  | 0.18088 | 0.18097 | 0.15848 | 0.18868 | --- | --- | 173.25 | 161.25 | 184.25 | 180.25 | --- | --- | 0.5056 | 0.47032 | 0.61371 | 0.50428 | --- | --- | 0.52348 | 0.062329 | 0.52085 | 1.1212 |  | -0.682 | -0.75434 | -0.48824 | -0.68461 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.2445 | 0.2321 | 0.2316 | 0.2218 |  |  | 163 | 160 | 159 | 155 |  |  | 0.20258 | 0.19018 | 0.18968 | 0.17988 | --- | --- | 147.25 | 144.25 | 143.25 | 139.25 | --- | --- | 0.38369 | 0.40038 | 0.39866 | 0.40864 | --- | --- | 0.39784 | 0.01039 | 0.39774 | 1.0266 |  | -0.95791 | -0.91533 | -0.91965 | -0.89492 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.257 | 0.2402 | 0.2335 | 0.2544 |  |  | 22 | 21 | 22 | 22 |  |  | 0.21508 | 0.19828 | 0.19158 | 0.21248 | --- | --- | 6.25 | 5.25 | 6.25 | 6.25 | --- | --- | 0.015339 | 0.013977 | 0.017221 | 0.015527 | --- | --- | 0.015516 | 0.0013301 | 0.015473 | 1.0892 |  | -4.1773 | -4.2704 | -4.0616 | -4.1652 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.2684 | 0.2515 | 0.2542 | 0.242 |  |  | 22 | 22 | 23 | 22 |  |  | 0.22648 | 0.20958 | 0.21227 | 0.20008 | --- | --- | 6.25 | 6.25 | 7.25 | 6.25 | --- | --- | 0.014567 | 0.015742 | 0.018028 | 0.016489 | --- | --- | 0.016207 | 0.0014494 | 0.016158 | 1.0932 |  | -4.229 | -4.1514 | -4.0158 | -4.105 | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.132 | 0.1352 | 0.1228 | 0.1351 |  |  | 141 | 143 | 144 | 150 |  |  | 0.090075 | 0.093275 | 0.080875 | 0.093175 | --- | --- | 125.25 | 127.25 | 128.25 | 134.25 | --- | --- | 0.73399 | 0.72012 | 0.83706 | 0.76055 | --- | --- | 0.76293 | 0.05219 | 0.76163 | 1.0692 |  | -0.30926 | -0.32833 | -0.17786 | -0.27371 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.1746 | 0.1662 | 0.159 | 0.1498 |  |  | 139 | 139 | 137 | 131 |  |  | 0.13268 | 0.12427 | 0.11708 | 0.10787 | --- | --- | 123.25 | 123.25 | 121.25 | 115.25 | --- | --- | 0.49036 | 0.5235 | 0.54668 | 0.56394 | --- | --- | 0.53112 | 0.031828 | 0.53039 | 1.0624 |  | -0.71262 | -0.64722 | -0.60389 | -0.5728 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.2264 | 0.2313 | 0.2184 | 0.2199 |  |  | 51 | 55 | 55 | 55 |  |  | 0.18448 | 0.18938 | 0.17648 | 0.17798 | --- | --- | 35.25 | 39.25 | 39.25 | 39.25 | --- | --- | 0.10086 | 0.1094 | 0.1174 | 0.11641 | --- | --- | 0.11102 | 0.0076494 | 0.11082 | 1.0727 |  | -2.294 | -2.2127 | -2.1422 | -2.1506 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.2232 | 0.2006 | 0.2104 | 0.1994 |  |  | 44 | 45 | 43 | 42 |  |  | 0.18128 | 0.15868 | 0.16848 | 0.15748 | --- | --- | 28.25 | 29.25 | 27.25 | 26.25 | --- | --- | 0.082261 | 0.097304 | 0.085378 | 0.08799 | --- | --- | 0.088233 | 0.0064849 | 0.08806 | 1.0747 |  | -2.4979 | -2.3299 | -2.4607 | -2.4305 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.258 | 0.2491 | 0.2469 | 0.2515 |  |  | 20 | 20 | 21 | 20 |  |  | 0.21608 | 0.20717 | 0.20498 | 0.20958 | --- | --- | 4.25 | 4.25 | 5.25 | 4.25 | --- | --- | 0.010382 | 0.010828 | 0.01352 | 0.010704 | --- | --- | 0.011359 | 0.0014529 | 0.011294 | 1.1289 |  | -4.5676 | -4.5256 | -4.3036 | -4.5371 | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.2904 | 0.26 | 0.2541 | 0.2482 |  |  | 20 | 22 | 21 | 20 |  |  | 0.24848 | 0.21808 | 0.21218 | 0.20628 | --- | --- | 4.25 | 6.25 | 5.25 | 4.25 | --- | --- | 0.0090286 | 0.015128 | 0.013061 | 0.010876 | --- | --- | 0.012023 | 0.0026459 | 0.011802 | 1.2509 |  | -4.7074 | -4.1912 | -4.3381 | -4.5212 | --- | ---
Hour 6:
Negative Control (Colony 1) | 0.3454 | 0.3352 | 0.34 | 0.3211 |  |  | 17 | 16 | 16 | 18 |  |  | 0.30347 | 0.29328 | 0.29808 | 0.27918 | --- | --- | 1.25 | 0.25 | 0.25 | 2.25 | --- | --- | 0.0021742 | 0.00044997 | 0.00044272 | 0.0042542 | --- | --- | 0.0018303 | 0.0018096 | 0.0011651 | 3.1309 |  | -6.1311 | -7.7063 | -7.7226 | -5.4598 | --- | ---
Negative Control (Colony 2) | 0.3626 | 0.3601 | 0.3534 | 0.3384 |  |  | 17 | 15 | 16 | 17 |  |  | 0.32067 | 0.31817 | 0.31148 | 0.29647 | --- | --- | 1.25 | -0.75 | 0.25 | 1.25 | --- | --- | 0.0020576 | -0.0012443 | 0.00042367 | 0.0022255 | --- | --- | 0.00086564 | 0.0016245 | #NUM! | #NUM! |  | -6.1862 | #NUM! | -7.7665 | -6.1078 | --- | ---
Positive Control (Colony 1) | 0.2813 | 0.2822 | 0.2814 | 0.2899 |  |  | 124 | 126 | 136 | 138 |  |  | 0.23938 | 0.24028 | 0.23947 | 0.24798 | --- | --- | 108.25 | 110.25 | 120.25 | 122.25 | --- | --- | 0.23871 | 0.24221 | 0.26506 | 0.26023 | --- | --- | 0.25155 | 0.013039 | 0.2513 | 1.0532 |  | -1.4325 | -1.418 | -1.3278 | -1.3462 | --- | ---
Positive Control (Colony 2) | 0.3011 | 0.3087 | 0.2982 | 0.2831 |  |  | 123 | 115 | 121 | 116 |  |  | 0.25917 | 0.26677 | 0.25628 | 0.24118 | --- | --- | 107.25 | 99.25 | 105.25 | 100.25 | --- | --- | 0.21843 | 0.19638 | 0.21679 | 0.21942 | --- | --- | 0.21275 | 0.010969 | 0.21253 | 1.0544 |  | -1.5213 | -1.6277 | -1.5288 | -1.5168 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 1) | 0.0712 | 0.0712 | 0.0723 | 0.0729 |  |  | 125 | 121 | 132 | 130 |  |  | 0.029275 | 0.029275 | 0.030375 | 0.030975 | --- | --- | 109.25 | 105.25 | 116.25 | 114.25 | --- | --- | 1.9699 | 1.8978 | 2.0202 | 1.947 | --- | --- | 1.9587 | 0.050851 | 1.9582 | 1.0263 |  | 0.67797 | 0.64067 | 0.70319 | 0.66628 | --- | ---
Test Device 1: J23101.BCD2.E0040.B0015 (Colony 2) | 0.0714 | 0.0674 | 0.0684 | 0.0647 |  |  | 88 | 87 | 90 | 83 |  |  | 0.029475 | 0.025475 | 0.026475 | 0.022775 | --- | --- | 72.25 | 71.25 | 74.25 | 67.25 | --- | --- | 1.2939 | 1.4763 | 1.4804 | 1.5587 | --- | --- | 1.4523 | 0.1122 | 1.4489 | 1.0827 |  | 0.25766 | 0.38956 | 0.3923 | 0.44382 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 1) | 0.2683 | 0.2648 | 0.2631 | 0.2861 |  |  | 177 | 175 | 182 | 190 |  |  | 0.22637 | 0.22287 | 0.22118 | 0.24418 | --- | --- | 161.25 | 159.25 | 166.25 | 174.25 | --- | --- | 0.376 | 0.37717 | 0.39677 | 0.37669 | --- | --- | 0.38166 | 0.010088 | 0.38156 | 1.0264 |  | -0.97817 | -0.97507 | -0.9244 | -0.97633 | --- | ---
Test Device 2: J23106.BCD2.E0040.B0015 (Colony 2) | 0.3512 | 0.345 | 0.3431 | 0.3215 |  |  | 163 | 170 | 174 | 160 |  |  | 0.30928 | 0.30307 | 0.30118 | 0.27958 | --- | --- | 147.25 | 154.25 | 158.25 | 144.25 | --- | --- | 0.25132 | 0.26865 | 0.27736 | 0.27235 | --- | --- | 0.26742 | 0.011311 | 0.26724 | 1.0439 |  | -1.381 | -1.3143 | -1.2824 | -1.3007 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 1) | 0.3073 | 0.3037 | 0.2895 | 0.3092 |  |  | 20 | 20 | 20 | 20 |  |  | 0.26538 | 0.26178 | 0.24757 | 0.26727 | --- | --- | 4.25 | 4.25 | 4.25 | 4.25 | --- | --- | 0.0084536 | 0.0085699 | 0.0090614 | 0.0083935 | --- | --- | 0.0086196 | 0.0003035 | 0.0086157 | 1.0353 |  | -4.7732 | -4.7595 | -4.7037 | -4.7803 | --- | ---
Test Device 3: J23117.BCD2.E0040.B0015 (Colony 2) | 0.3716 | 0.3474 | 0.3446 | 0.3314 |  |  | 19 | 20 | 19 | 19 |  |  | 0.32967 | 0.30547 | 0.30268 | 0.28947 | --- | --- | 3.25 | 4.25 | 3.25 | 3.25 | --- | --- | 0.0052037 | 0.0073439 | 0.0056679 | 0.0059263 | --- | --- | 0.0060355 | 0.00092212 | 0.0059857 | 1.1579 |  | -5.2584 | -4.9139 | -5.1729 | -5.1283 | --- | ---
Test Device 4: J23101+I13504 (Colony 1) | 0.2284 | 0.2231 | 0.2077 | 0.2317 |  |  | 128 | 133 | 138 | 141 |  |  | 0.18648 | 0.18118 | 0.16578 | 0.18977 | --- | --- | 112.25 | 117.25 | 122.25 | 125.25 | --- | --- | 0.31775 | 0.34161 | 0.38926 | 0.34838 | --- | --- | 0.34925 | 0.029736 | 0.34832 | 1.0876 |  | -1.1465 | -1.0741 | -0.9435 | -1.0545 | --- | ---
Test Device 4: J23101+I13504 (Colony 2) | 0.3092 | 0.3063 | 0.2912 | 0.2736 |  |  | 142 | 138 | 142 | 139 |  |  | 0.26727 | 0.26438 | 0.24928 | 0.23168 | --- | --- | 126.25 | 122.25 | 126.25 | 123.25 | --- | --- | 0.24934 | 0.24409 | 0.26734 | 0.28082 | --- | --- | 0.2604 | 0.016868 | 0.25999 | 1.0665 |  | -1.3889 | -1.4102 | -1.3192 | -1.2701 | --- | ---
Test Device 5: J23106+I13504 (Colony 1) | 0.301 | 0.3001 | 0.2829 | 0.3033 |  |  | 42 | 44 | 46 | 46 |  |  | 0.25907 | 0.25817 | 0.24097 | 0.26138 | --- | --- | 26.25 | 28.25 | 30.25 | 30.25 | --- | --- | 0.053483 | 0.057759 | 0.066263 | 0.061091 | --- | --- | 0.059649 | 0.0053977 | 0.059466 | 1.0946 |  | -2.9284 | -2.8515 | -2.7141 | -2.7954 | --- | ---
Test Device 5: J23106+I13504 (Colony 2) | 0.2743 | 0.2718 | 0.2699 | 0.2667 |  |  | 37 | 35 | 35 | 34 |  |  | 0.23237 | 0.22987 | 0.22797 | 0.22478 | --- | --- | 21.25 | 19.25 | 19.25 | 18.25 | --- | --- | 0.048271 | 0.044203 | 0.044572 | 0.042858 | --- | --- | 0.044976 | 0.0023168 | 0.044932 | 1.052 |  | -3.0309 | -3.119 | -3.1107 | -3.1499 | --- | ---
Test Device 6: J23117+I13504 (Colony 1) | 0.3262 | 0.3275 | 0.3065 | 0.3301 |  |  | 17 | 19 | 19 | 18 |  |  | 0.28427 | 0.28558 | 0.26458 | 0.28818 | --- | --- | 1.25 | 3.25 | 3.25 | 2.25 | --- | --- | 0.0023211 | 0.0060073 | 0.0064841 | 0.0041214 | --- | --- | 0.0047334 | 0.0019045 | 0.0043935 | 1.5988 |  | -6.0657 | -5.1148 | -5.0384 | -5.4916 | --- | ---
Test Device 6: J23117+I13504 (Colony 2) | 0.4117 | 0.3814 | 0.3921 | 0.3602 |  |  | 17 | 17 | 17 | 18 |  |  | 0.36978 | 0.33948 | 0.35018 | 0.31828 | --- | --- | 1.25 | 1.25 | 1.25 | 2.25 | --- | --- | 0.0017844 | 0.0019436 | 0.0018843 | 0.0037316 | --- | --- | 0.002336 | 0.00093274 | 0.0022222 | 1.4153 |  | -6.3287 | -6.2432 | -6.2742 | -5.5909 | --- | ---

## Calculate Target Dilution
target Abs600 |  | 0.02
target volume (mL) |  | 12
sample | Abs600 Reading | Volume of Preloading Culture | Volume of Preloading Media
positive control | 0.6152 | 0.41754 | 11.5825 | Enter fluorescence measurements into blue cells
negative control | 0.6728 | 0.37951 | 11.6205 | Gold cells are calculated
device 1 | 0.725 | 0.35057 | 11.6494
device 2 | 0.5333 | 0.48691 | 11.5131
device 3 | 0.5665 | 0.45619 | 11.5438
device 4 | 0.5548 | 0.46656 | 11.5334
device 5 | 0.5826 | 0.44264 | 11.5574
device 6 | 0.3254 | 0.84211 | 11.1579
media+chl | 0.0404

## Plate Reader Settings
SAFIRE II; Serial number: 501000006; Firmware: V 1.35 08/2005 Safire2; XFLUOR4SAFIREII Version: V 4.62n
Fluorescence Settings |  |  | Absorbance Settings
Measurement mode: | Fluorescence Bottom |  | Measurement mode: | Absorbance
Excitation wavelength: | 485 | nm | Measurement wavelength: | 600
Emission wavelength: | 530 | nm | Number of reads: | 3
Excitation bandwidth: | 5 | nm | Plate definition file: | COS96fb.pdf | nm
Emission bandwidth: | 5 | nm
Gain (Manual): | 50
Number of reads: | 6
FlashMode: | High sensitivity
Integration time: | 40 | µs
Lag time: | 0 | µs
Plate definition file: | COS96fb.pdf
Time between move and flash: | 10 | ms